## AU400
AU400
No. | Short name | ADD | Ds
S57 | Cuv. wheel sensor Stop Position | 43 | DS2
S58 | Cuv. wheel sensor Origin | 42 | DS1
S52 | Reagent fridge sensor Bottle detection | 44 | DS5
S51 | Reagent fridge sensor Lid open | 44 | DS4
S49 | Reagent fridge sensor Outside stop position | 44 | DS3
S50 | Reagent fridge sensor Inside stop position | 44 | DS2
S48 | Reagent fridge sensor Origin | 44 | DS1
S31 | Cooling sampler sensor Cup detection | 45 | DS5
S32 | Cooling sampler sensor Lid open | 45 | DS4
S30 | Cooling sampler sensor Stop position | 45 | DS2
S29 | Cooling sampler sensor Origin | 45 | DS1
S62 | Cuv. wash unit sensor Lower point position | 46 | DS2
S61 | Cuv. wash unit sensor Upper point position | 46 | DS1
S44 | Reagent probe sensor Normal stop position | 47 | DS6
S43 | Reagent probe sensor Rotation origin | 47 | DS5
S45 | Reagent probe sensor Special stop position (reagent) | 47 | DS4
S55 | Reagent probe sensor Level detection | 47 | DS2
S46 | Reagent probe sensor Upper point position | 47 | DS1
S35 | Sample probe sensor Normal stop position | 48 | DS6
S33 | Sample probe sensor Rotation origin | 48 | DS5
S34 | Sample probe sensor Special stop position(Cuv.) | 48 | DS4
S40 | Sample probe sensor Level detection | 48 | DS2
S36 | Sample probe sensor Upper point position | 48 | DS1
S7 | Rack transfer unit Rack detection | 49 | DS6
S3 | Rack transfer unit Rack presence | 49 | DS5
S1 | Rack transfer unit Rack feed pitch | 49 | DS4
S2 | Rack transfer unit Rack receive pitch | 49 | DS3
S28 | Rack transfer unit Re-run rack arrival | 49 | DS2
S27 | Rack transfer unit Re-run rack pass | 49 | DS1
S4 | Rack transfer unit Rack identification 1 | 4A | DS7
S5 | Rack transfer unit Rack identification 2 | 4A | DS6
S6 | Rack transfer unit Rack identification 1 | 4A | DS5
S8 | Rack transfer unit Sample cup detec tion | 4A | DS4
S16 | Rack transfer unit Rack stepping | 4A | DS3
S10 | Rack transfer unit ID mirror stop position | 4C | DS2
S9 | Rack transfer unit ID mirror origin | 4C | DS1
S14 | Rack transfer unit Claw feed origin 2 | 4D | DS2
S11 | Rack transfer unit Claw feed origin 1 | 4D | DS1
S20 | Rack transfer unit Re-run rack detec tion | 4E | DS6
S19 | Rack transfer unit Re-run rack feed pitch | 4E | DS4
S18 | Rack transfer unit Re-run rack receive pitch | 4E | DS3
S17 | Rack transfer unit Rack arrival | 4E | DS2
S21 | Rack transfer unit Storage rack presence | 4E | DS2
S26 | Rack transfer unit Rack full | 51 | DS6
S25 | Rack transfer unit Push out position | 51 | DS5
S24 | Rack transfer unit Standby position | 51 | DS4
S39 | Dispenser sensor Sample lower point | 53 | DS6
S38 | Dispenser sensor Sample upper point | 53 | DS5
S54 | Dispenser sensor Reagent lower point | 53 | DS2
S53 | Dispenser sensor Reagent upper point | 53 | DS1
S59 | Mix unit sensor Stop position | 55 | DS5
S58 | Mix unit sensor Origin | 55 | DS4
S60 | Mix unit sensor Upper point position | 55 | DS1
S47 | Sample probe sensor Crash detection | 57 | DS7
S37 | Reagent probe sensor Reagent probe sensor | 58 | DS8

## AU480
AU480
No. | Short name | ADD | Ds
DAS06 | DA SENSOR Outer stop pos | 02 | DS8
DAS02 | DA SENSOR Inner stop pos | 02 | DS7
DAS06 | DA SENSOR R Bottle detection | 02 | DS5
DAS03 | DA SENSOR Cover open | 02 | DS4
DAS01 | DA SENSOR Rotation origin | 02 | DS1
DCS02 | DC SENSOR Stop pos. | 05 | SD7
DCS05 | DC SENSOR Sample detection | 05 | DS4
DCS04 | DC SENSOR Small cover open | 05 | DS3
DCS03 | DC SENSOR Large cover open | 05 | DS2
DCS01 | DC SENSOR Rotation origin | 05 | DS1
FAS01 | FA SENSOR Liquid level | 09 | DS4
FAS07 | FA SENSOR Crash | 09 | DS3
FAS01 | FA SENSOR Upper pos | 09 | DS2
FAS05 | FA SENSOR Ratotion stop pos. 1 | 0D | DS3
FAS04 | FA SENSOR Stop pos | 0D | DS2
FAS03 | FA SENSOR Rotation origin | 0D | DS1
FBS02 | FB SENSOR Lower position | 10 | DS2
FBS01 | FB SENSOR Upper position | 10 | DS1
FCS05 | FC SENSOR Lower position | 13 | DS2
FCS01 | FC SENSOR Upper position | 13 | DS1
FCS03 | FC SENSOR Stop position | 16 | DS2
FCS01 | FC SENSOR Rotation origin | 16 | DS1
FCS04 | FC SENSOR Rotation detection | 19 | DS7
FDS08 | FD SENSOR Liquid level detection | 1C | DS4
FDS07 | FD SENSOR Crush | 1C | DS3
FDS01 | FD SENSOR Upper position | 1C | DS2
FDS05 | FD SENSOR Rotation stop pos. 1 | 21 | DS3
FDS04 | FD SENSOR Stop pos. | 21 | DS2
FDS03 | FD SENSOR Rotation origin | 21 | DS1
V29 | Sample aspiration valve | 24 | DS8
V26 | Liquid dripping aspiration valve | 24 | DS7
V23 | Aspiration valve for OVF | 24 | DS6
V25 | Air release valve for dripping liquid | 24 | DS5
V31 | Air release valve for aspiration | 24 | DS4
V22 | Air release valve for OVF | 24 | DS3
V28 | Air release valve for sample | 24 | DS2
V32 | Wash water aspiration valve | 24 | DS1
V2 | DI water supply valve | 25 | DS8
V1 | OPD valve | 25 | DS7
P4 | Vacuum pump | 25 | DS5
V34 | Vacuum tank waste water valve | 25 | DS2
P9 | Degasser | 25 | DS1
V19 | Wash water dripping valve | 26 | DS8
V21 | Detergent dripping valve | 26 | DS7
V16 | Wash water dispense valve | 26 | DS6
V18 | Detergent dispense valve | 26 | DS5
V11 | Mix wash water valv | 27 | DS8
V9 | Interior wash back valve | 27 | DS7
V6 | R inner wash valve | 27 | DS6
V5 | S inner wash valve | 27 | DS5
V30 | Sample drain valve | 27 | DS4
V27 | Dripping liquid drain valve | 27 | DS3
V33 | Aspiration drain valve | 27 | DS2
V24 | OVF drain valve | 27 | DS1
| Test counter | 28 | DS4
| Detergent dilution valve | 28 | DS3
| Mix detergent valve | 28 | DS2
| Probe outer wash | 28 | DS1
| DS rack stopper | 29 | DS5
P3 | Diluted detergent supply | 29 | DS2
P1 | DI water | 29 | DS1
| DIAG LED | 2B | DS3
| STAT LED | 2B | DS1
| Diluted detergent Lower limit | 2D | DS8
| Diluted detergent Lower position | 2D | DS7
| Diluted detergent Upper pos. | 2D | DS6
| Diluted detergent Upper limit | 2D | DS5
| DI water Lower limit | 2D | DS4
| DI water Lower pos. | 2D | DS3
| DI water Upper pos. | 2D | DS2
| DI water Upper limit | 2D | DS1
AAS04 | Vacuum tank Pressure | 03 | DS1
AAS07 | Rear DIAG SW | 31 | DS5
AAS06 | DIAG SW | 31 | DS2
| DI water supply valve status | 31 | DS8
| OPD valve status | 31 | DS7
SAS02 | S SA lower limit | 34 | DS2
SAS01 | S SA upper Limit | 34 | DS1
| CLOT DETECTION SENSOR Foreign object. Aspiration 1 | 35 | DS8
| CLOT DETECTION SENSOR Foreign object. Aspiration 2 | 35 | DS7
| CLOT DETECTION SENSOR Foreign object. Aspiration 3 | 35 | DS6
| CLOT DETECTION SENSOR Air bubble incorporation | 35 | DS5
| Air aspiration | 35 | DS4
| Invalid Command | 35 | DS3
| CLOT DETECTION SENSOR Communication error | 36 | DS8
SAS42 | SA SENSOR R SA lower limit | 38 | DS2
SAS41 | SA SENSOR R SA upper limit | 38 | DS1
SAS42 | SA SENSOR ISE SA lower limit | 3A | DS2
SAS11 | SA SENSOR ISE SA upper limit | 3A | DS1
CBS20 | CB SENSOR Cover open | 3C | DS6
CBS01 | CB SENSOR Rack detection | 3C | DS5
CBS14 | CB SENSOR Rack step 10 | 3C | DS4
CBS13 | CB SENSOR Rack step 9 | 3C | DS3
CBS12 | CB SENSOR Rack step 8 | 3C | DS2
CBS11 | CB SENSOR Rack step 7 | 3C | DS1
CBS09 | CB SENSOR Rack step 6 | 3E | DS8
CBS08 | CB SENSOR Rack step 5 | 3E | DS7
CBS07 | CB SENSOR Rack step 4 | 3E | DS6
CBS06 | CB SENSOR Rack step 3 | 3E | DS5
CBS05 | CB SENSOR Rack step 2 | 3E | DS4
CBS04 | CB SENSOR Rack step 1 | 3E | DS3
CBS03 | Mirror stop | 3E | DS2
CBS02 | Mirror Origin | 3E | DS1
CGS01 | CG SENSOR Sample feed origin | 40 | DS1
CAS06 | CA SENSOR Cover open | 42 | DS8
CAS05-3 | CA SENSOR Rack type 2 | 42 | DS7
CAS05-2 | CA SENSOR Rack type 1 | 42 | DS6
CAS05-1 | CA SENSOR Rack type 0 | 42 | DS5
CAS04 | CA SENSOR Rack detection | 42 | DS4
CAS03 | CA SENSOR Re-run rack arrival | 42 | DS3
CAS02 | CA SENSOR Rack feed pitch | 42 | DS2
CAS01 | CA SENSOR Rack receptacle pitch | 42 | DS1
CCS03 | CC SENSOR Rack feed pitch | 45 | DS3
CCS02 | CC SENSOR Rack receptacle pitch | 45 | DS2
CCS01 | CC SENSOR Rack arrival | 45 | DS1
CDS03 | CD SENSOR Re-run rack pass | 47 | DS3
CDS02 | CD SENSOR Rack storage present | 47 | DS2
CDS01 | CD SENSOR Re-run rack detection | 47 | DS1
CBS19 | CB, CG SENSOR Cup detect 5 | 4A | DS8
CBS18 | CB, CG SENSOR Cup detect 4 | 4A | DS7
CBS17 | CB, CG SENSOR Cup detect 3 | 4A | DS6
CBS16 | CB, CG SENSOR Cup detect 2 | 4A | DS5
CBS15 | CB, CG SENSOR Cup detect 1 | 4A | DS4
CGS01 | CB, CG SENSOR ID feed origin | 4A | DS1
CHS01 | CE, CH SENSOR Rack full | 4C | DS3
CES02 | CE, CH SENSOR Feeding position | 4C | DS2
CES01 | CE, CH SENSOR Waiting Position | 4C | DS1
IDS02 | ISE SENSOR MID solution present | 4F | DS3
IDS01 | ISE SENSOR Buffer solution present | 4F | DS2
IAS01 | ISE SENSOR Cover open/close | 4F | DS1
IDS03 | ISE SENSOR REF solution present | 52 | DS2
IAS02 | ISE SENSOR Pot over flow detection | 5C | DS8
IAS03 | ISE SENSOR Analog power ON | 5C | DS6
IAS04 | ISE SENSOR Cable disconnection | 5C | DS5
GAS04 | GA SENSOR Rotations not permitted | 62 | DS3
GAS02 | GA SENSOR Stop pos. | 62 | DS2
GAS01 | GA SENSOR Rotation origin | 62 | DS1

## AU640
AU640
No. | Short name | ADD | Ds
C1 | S counter | 14 | DS4
P1 | DI water pump | 14 | DS7
P2 | Detergent dilution pump | 14 | DS8
P4 | Detergent A pump | 53 | DS8
P5 | Detergent B pump | 53 | DS7
P6 | Vacuum pump | 15 | DS1
P12 | Pure detergent A pump | 14 | DS5
P13 | Pure detergent B pump | 14 | DS6
S63 | Vacuum tank full | 0C | DS7
S64 | Vacuum tank pressure SW | 0C | DS8
S65 | DI water (overflow) | 0C | DS2
S68 | DI water (empty) | 0C | DS1
S69 | Pure detergent A ( empty) | 0C | DS5
S70 | Pure detergent B ( empty) | 0C | DS6
S71 | Detergent A (overflow) | 0D | DS4
S72 | Detergent A (upper) | 0D | DS3
S73 | Detergent A (lower) | 0D | DS2
S74 | Detergent A (empty) | 0D | DS1
S75 | Detergent B (overflow) | 0D | DS8
S76 | Detergent B (upper) | 0D | DS7
S77 | Detergent B (lower) | 0D | DS6
S78 | Detergent B (empty) | 0D | DS5
SW1 | STAT SW | 0A | DS4
SW2 | DIAG SW | 0A | DS1
V3 | Detergent A dilution valve | 10 | DS3
V4 | Detergent B dilution valve | 10 | DS4
V5 | S inner wash valve | 10 | DS5
V6 | R1 inner wash valve | 10 | DS6
V7 | R2 inner wash valve | 10 | DS7
V10 | Probe outer wash valve | 10 | DS8
V11 | S/R1 M1 wash valve | 11 | DS1
V12 | R2 M2 wash valve | 11 | DS6
V16 | Wash water disp. valve | 12 | DS3
V17 | Detergent A disp. valve | 12 | DS1
V18 | Detergent B disp. valve | 12 | DS2
V19 | Wash water drip valve | 12 | DS6
V20 | Detergent A drip valve | 12 | DS4
V21 | Detergent B drip valve | 12 | DS5
V22 | OVF air release valve | 12 | DS8
V23 | OVF asp. valve | 13 | DS3
V24 | OVF drain valve | 13 | DS7
V25 | Drip air release valve | 13 | DS2
V26 | Drip asp. valve | 13 | DS4
V27 | Liquid dripping drain valve | 14 | DS1
V28 | Conc. waste air release valve | 12 | DS7
V29 | Mixture asp. valve | 14 | DS2
V30 | Mixture drain valve | 13 | DS6
V31 | Aspiration air release valve | 13 | DS1
V32 | Wash water asp. valve | 13 | DS5
V33 | Aspiration drain valve | 13 | DS8
V34 | Vacuum waste valve | 14 | DS3
V40 | Air drive valve | 15 | DS8
V41 | S/R1 M1detergent B valve | 11 | DS3
V45 | R2 M2 detergent B valve | 11 | DS8
M1 | CA rack feed | 4C | DS7
S1 | CA feed pitch | 4C | DS4
S2 | CA receptacle pitch | 4C | DS3
S3 | CA rack present | 4C | DS5
S4 | Rack type 1 | 4D | DS7
S5 | Rack type 2 | 4D | DS6
S6 | Rack type 3 | 4D | DS5
S91 | LA rack arrival | 4E | DS4
S92 | LA manual rack set | 4E | DS3
S93 | LA rack pass | 4E | DS1
S112 | CA repeat run arrival | 4C | DS2
M2 | CB belt | 4C | DS8
S7 | CB rack detection | 4C | DS6
S8 | CB cup detection | 4D | DS4
S16 | Rack feed detection | 4E | DS8
M6 | CC rack send | 51 | DS8
S100 | CC cover open | 51 | DS6
S101 | CC rack arrival | 52 | DS1
S102 | CC pre-rack detection | 51 | DS2
S103 | CC rack feed origin | 51 | DS4
S104 | CC rack feed pos. | 51 | DS1
S105 | CC shaft lower pos. | 51 | DS5
S106 | CC rack push pos. | 51 | DS3
M7 | CD rack send | 51 | DS7
S107 |  | 53 | DS1
S109 | CD retrieval pos. | 53 | DS3
S110 | CD rack repeat pos. | 53 | DS2
S111 | CD rack arrival detection | 52 | DS2
S113 | CD repeat run rack pass | 4C | DS1
S114 | CD retrieval rack pass | 52 | DS3
M8 | CE rack push | 52 | DS8
S24 | CE rack wait pos. | 52 | DS4
S25 | CE push Pos. | 52 | DS5
S11 | ID pawl origin | 50 | DS1
S14 | S pawl origin | 50 | DS2
M4 | ID pawl | 50 | DS7
M5 | S pawl | 50 | DS8
S26 | Rack full | 52 | DS6
M16 | R1 tray rot. | 44 | DS8
S48 | R1 tray rot. origin | 44 | DS1
S49 | R1 tray stop pos. (33) | 44 | DS2
S50 | R1 tray stop pos. (48) | 44 | DS3
S51 | R1 cover open | 44 | DS4
S52 | R1 bottle detection | 44 | DS5
M28 | R2 tray rot. | 45 | DS8
S120 | R2 tray rot. origin | 45 | DS1
S121 | R2 tray stop pos. (33) | 45 | DS2
S122 | R2 tray stop pos. (48) | 45 | DS3
S123 | R2 cover open | 45 | DS4
S124 | R2 bottle detection | 45 | DS5
M9 | STAT rot. | 4F | DS7
S29 | STAT rot. origin | 4F | DS1
S30 | STAT rot. stop pos. | 4F | DS2
S31 | STAT sample tube exist | 4F | DS3
S32 | STAT cover open | 4F | DS4
M34 | Probe inner wash | 58 | DS8
S134 | EC upper pos. | 58 | DS1
S135 | EC lower pos. | 58 | DS2
M10 | S probe rot. | 4A | DS8
S33 | S rot. origin | 4A | DS5
S34 | S rot cuv pos. | 4A | DS4
S35 | S rot. stop pos. | 4A | DS6
M11 | S probe U/D | 4A | DS7
S36 | S probe middle pos. | 4B | DS1
S37 | S probe crash detection | 4A | DS3
S40 | S liquid level detection | 4A | DS2
S42 | Sample probe upper pos. | 4A | DS1
M22 | Cuv wash U/D | 46 | DS7
S61 | FB upper pos. | 46 | DS1
S62 | FB down detection | 46 | DS2
M21 | M1 mix | 58 | DS7
M19 | M1 rot. | 57 | DS8
S58 | M1 rot. origin | 57 | DS5
S59 | M1 rot. stop pos. | 57 | DS6
M20 | M1 U/D | 57 | DS7
S60 | M1 upper pos. | 57 | DS1
M32 | M2 mix | 49 | DS7
M30 | M2 rot. | 48 | DS8
S130 | M2 rot. origin | 48 | DS5
S131 | M2 rot. stop pos. | 48 | DS6
M31 | M2 U/D | 48 | DS7
S132 | M2 upper pos. | 48 | DS1
M14 | R1 probe rot. | 47 | DS8
S43 | R1 probe rot. origin | 47 | DS5
S44 | R1 probe rot. cuv pos. | 47 | DS6
S45 | R1 probe rot. stop pos. | 47 | DS4
M15 | R1 probe U/D | 47 | DS7
S46 | R1 probe upper pos. | 47 | DS1
S47 | R1 probe crash detection | 47 | DS3
S55 | R1 liquid level detection | 47 | DS2
M26 | R2 probe rot. | 59 | DS8
S115 | R2 probe rot. origin | 59 | DS5
S116 | R2 probe rot. cuv pos. | 59 | DS6
S117 | R2 probe rot. stop pos. | 59 | DS4
M27 | R2 probe U/D | 59 | DS7
S118 | R2 probe upper pos. | 59 | DS1
S119 | R2 probe crash detection | 59 | DS3
S127 | R2 liquid level detection | 59 | DS2
M18 | Cuv rot. | 40 | DS3
S56 | Cuv wheel origin | 43 | DS1
S57 | Cuv wheel stop pos. | 43 | DS2
P3 | Degassor | 15 | DS2
S86 | ISE B empty | 0E | DS3
S87 | ISE MID empty | 0E | DS2
S88 | ISE REF empty | 0E | DS1
S79 | OPH1 full | 0C | DS4
S81 | OPH2 full | 0C | DS3
M12 | S disp. | 55 | DS4
M17 | R1 disp. | 54 | DS4
M29 | R2 disp. | 54 | DS8
S125 | R2 disp. upper pos. | 54 | DS5
S126 | R2 disp. lower pos. | 54 | DS6
S38 | S disp. upper pos. | 55 | DS1
S39 | S disp. lower pos. | 55 | DS2
S41 | S clog detection | 56 | DS1
S53 | R1 disp. upper pos. | 54 | DS1
S54 | R1 disp. lower pos. | 54 | DS2

## AU680
AU680
No. | Short name | ADD | Ds
CAS06,CAS07 | CB SENSOR CB Cover | 02 | DS5
CBS01 | CB SENSOR CB Stop | 02 | DS4
CBS03 | CB SENSOR Rack detect 2 | 02 | DS2
CBS06-3 | Rack type Front Side | 03 | DS8
CBS06-2 | Rack type Rack type | 03 | DS7
CBS06-1 | Rack type Rack ID Rabel side | 03 | DS6
CBS04 | Rack detect 3 | 03 | DS3
CBS02 | Rack detect 1 | 03 | DS1
CDS14 | CC, CD SENSOR Rack step 6 | 07 | DS8
CDS13 | CC, CD SENSOR Rack step 5 | 07 | DS7
CDS12 | CC, CD SENSOR Rack step 4 | 07 | DS6
CDS11 | CC, CD SENSOR Rack step 3 | 07 | DS5
CDS10 | CC, CD SENSOR Rack step 2 | 07 | DS4
CDS9 | CC, CD SENSOR Rack step1 | 07 | DS3
CCS02 | CC, CD SENSOR CC Stop | 07 | DS2
CCS01 | CC, CD SENSOR CC Origin | 07 | DS1
CDS26 | CC, CD SENSOR CD Rack check | 0A | DS8
CDS18 | CC, CD SENSOR Rack step 10 | 0A | DS6
CDS17 | CC, CD SENSOR Rack step 9 | 0A | DS5
CDS16 | CC, CD SENSOR Rack step 8 | 0A | DS4
CDS15 | CC, CD SENSOR Rack step 7 | 0A | DS3
CCS04 | CC, CD SENSOR Shutter stop | 0A | DS2
CCS03 | CC, CD SENSOR Shutter origin | 0A | DS1
CDS22 | CD SENSOR Cup detect 3 | 0C | DS7
CDS21 | CD SENSOR Cup detect 2 | 0C | DS6
CDS20 | CD SENSOR Cup detect 1 | 0C | DS5
CDS04 | CD SENSOR Lever1 back | 0C | DS4
CDS03 | CD SENSOR Lever1 forward | 0C | DS3
CDS02 | CD SENSOR Lever1 stop | 0C | DS2
CDS01 | CD SENSOR Lever1 origin | 0C | DS1
CDS24 | CD SENSOR Cup detect 5 | 0F | DS8
CDS23 | CD SENSOR Cup detect 4 | 0F | DS7
CDS19 | CD SENSOR CD Rack pass 1 | 0F | DS5
CDS08 | CD SENSOR Lever2 back | 0F | DS4
CDS07 | CD SENSOR Lever2 forward | 0F | DS3
CDS06 | CD SENSOR Lever2 stop | 0F | DS2
CDS05 | CD SENSOR Lever2 origin | 0F | DS1
CES11 | CE SENSOR STATUS CE Rack pass 2 | 12 | DS8
CES09 | CE SENSOR STATUS CE Rack pass 1 | 12 | DS7
CES07 | CE SENSOR STATUS CE Rack detect 2 | 12 | DS6
CES06 | CE SENSOR STATUS CE Rack detect 1 | 12 | DS5
CES02 | CE SENSOR STATUS CE Lever push | 12 | DS4
CES03 | CE SENSOR STATUS Front rack detect | 12 | DS3
CES01 | CE SENSOR STATUS Lever origin | 12 | DS1
CES05 | CE SENSOR Rack support stop | 14 | DS2
CES04 | CE SENSOR Rack support origin | 14 | DS1
CAS01 | Rack full detect | 15 | DS3
CAS03 | CA, CF SENSOR CA Rack pass | 17 | DS5
CAS04 | CA, CF SENSOR CD Cover | 17 | DS3
CFS02 | CA, CF SENSOR CF Lever1 stop | 17 | DS2
CFS01 | CA, CF SENSOR CF Lever1 origin | 17 | DS1
CFS04 | CF LEVER2 SENSOR Lever2 stop | 1A | DS2
CFS03 | CF LEVER2 SENSOR Lever2 origin | 1A | DS1
CGS02 | CG SENSOR Lever stop | 1C | DS2
CGS01 | CG SENSOR Lever origin | 1C | DS1
CHS02 | CH SENSOR Lever stop | 1E | DS8
CAS05 | CH SENSOR CE Cover | 1E | DS3
CAS05 | CH SENSOR Rack detect | 1E | DS2
CHS01 | CH SENSOR Lever origin | 1E | DS1
CJS02 | CJ SENSOR CJ Stop | 22 | DS2
CJS01 | CJ SENSOR CJ Origin | 22 | DS1
DAS02 | DA01 SENSOR R1 Stop | 24 | DS7
DAS04 | DA01 SENSOR R1 Bottle detect | 24 | DS3
DAS03 | DA01 SENSOR R1 Cover | 24 | DS2
DAS01 | DA01 SENSOR R1 Origin | 24 | DS1
DAS12 | DA11 SENSOR R2 Stop | 28 | DS7
DAS14 | DA11 SENSOR R2 Bottle detect | 28 | DS3
DAS13 | DA11 SENSOR R2 Cover | 28 | DS2
DAS11 | DA11 SENSOR R2 Origin | 28 | DS1
DCS02 | DC SENSOR DC Stop | 2C | DS7
DCS05 | DC SENSOR Cup detect | 2C | DS4
DCS04 | DC SENSOR DC Cover 2 | 2C | DS3
DCS03 | DC SENSOR DC Cover 1 | 2C | DS2
DCS01 | DC SENSOR DC Origin | 2C | DS1
| FA UPPER/LOWER POS. SENSOR Liquid level | 30 | DS4
FAS07 | FA UPPER/LOWER POS. SENSOR S Crash | 30 | DS3
FAS01 | FA UPPER/LOWER POS. SENSOR S Upper | 30 | DS2
FAS02 | FA UPPER/LOWER POS. SENSOR S Middle | 30 | DS1
FAS06 | FA ROTATION SENSOR S Rotation pos.2 | 33 | DS4
FAS05 | FA ROTATION SENSOR S Rotation pos.1 | 33 | DS3
FAS04 | FA ROTATION SENSOR S Rotation stop | 33 | DS2
FAS03 | FA ROTATION SENSOR S Rotation origin | 33 | DS1
FBS02 | FB SENSOR Lower pos. | 35 | DS2
FBS01 | FB SENSOR Upper pos. | 35 | DS1
FCS05 | FC01UPPER/LOWER POS. SENSOR Lower pos. | 37 | DS2
FCS01 | FC01UPPER/LOWER POS. SENSOR Upper pos. | 37 | DS1
FCS03 | FC01 ROTATION SENSOR Rotation stop | 39 | DS2
FCS02 | FC01 ROTATION SENSOR Rotation stop | 39 | DS1
FCS04 | FC01 MIX SENSOR FC01 Mix monitor | 3B | DS7
FCS15 | FC11 UPPER/LOWER POS. SENSOR Lower pos. | 3E | DS2
FCS11 | FC11 UPPER/LOWER POS. SENSOR Upper pos. | 3E | DS1
FCS13 | FC11 ROTATION SENSOR Rotation stop | 40 | DS2
FCS12 | FC11 ROTATION SENSOR Rotation origin | 40 | DS1
FCS14 | FC11 Mix monitor | 42 | DS7
FDS07 | FD01 UPPER/LOWER POS. SENSOR R1 Crash | 45 | DS3
FDS01 | FD01 UPPER/LOWER POS. SENSOR R1 Upper pos. | 45 | DS2
| FD01 UPPER/LOWER POS. SENSOR R1 Liquid level | 45 | DS4
FDS06 | FD01 ROTATION SENSOR R1 Rotation pos.2 | 48 | DS4
FDS05 | FD01 ROTATION SENSOR R1 Rotation pos.1 | 48 | DS3
FDS04 | FD01 ROTATION SENSOR R1 Rotation stop | 48 | DS2
FDS03 | FD01 ROTATION SENSOR R1 Rotation origin | 48 | DS1
| FD11 UPPER/LOWER POS. SENSOR R2 Liquid level | 4A | DS4
FDS17 | FD11 UPPER/LOWER POS. SENSOR R2 Crash | 4A | DS3
FDS11 | FD11 UPPER/LOWER POS. SENSOR R2 Upper pos. | 4A | DS2
FDS16 | FD11 ROTATION SENSOR R2 Rotation pos.2 | 4D | DS4
FDS15 | FD11 ROTATION SENSOR R2 Rotation pos.1 | 4D | DS3
FDS14 | FD11 ROTATION SENSOR R2 Rotation stop | 4D | DS2
FDS13 | FD11 ROTATION SENSOR R2 Rotation origin | 4D | DS1
| S DISPENSER SENSOR S Dispensing pressure | 4F | DS3
SAS02 | S DISPENSER SENSOR S Dispenser lower pos. | 4F | DS2
SAS01 | S DISPENSER SENSOR S Dispenser upper pos. | 4F | DS1
SAS12 | R1 DISPENSER SENSOR R1 Dispenser lower pos. | 52 | DS2
SAS11 | R1 DISPENSER SENSOR R1 Dispenser upper pos. | 52 | DS1
SAS22 | R2 DISPENSER SENSOR R2 Dispenser lower pos. | 54 | DS2
SAS21 | R2 DISPENSER SENSOR R2 Dispenser upper pos. | 54 | DS1
SAS32 | WASH DISPENSER SENSOR Wash dispenser lower pos. | 56 | DS2
SAS31 | WASH DISPENSER SENSOR Wash dispenser upper pos. | 56 | DS1
ISES02 | ISE SENSOR Mid liquid level | 5A | DS3
ISES01 | ISE SENSOR Buffer liquid level | 5A | DS2
ISES05 | ISE SENSOR ISE Cover | 5A | DS1
ISES03 | BUFFER LIQUID DISPENSER SENSOR Ref liquid level | 5D | DS3
SAS42 | BUFFER LIQUID DISPENSER SENSOR B. dispenser lower pos. | 5D | DS2
SAS41 | BUFFER LIQUID DISPENSER SENSOR B. dispenser upper pos. | 5D | DS1
ISES04 | ISE SENSOR ISE Sample pot liquid level | 5E | DS9
| ISE SENSOR ISE AP ON | 5E | DS6
| ISE SENSOR ISE Sample Pot Liquid Level Sensor disconnected | 5E | DS5
| GA SENSOR Rotations not permitted | 64 | DS3
GAS02 | GA SENSOR GA Stop | 64 | DS2
GAS01 | GA SENSOR GA Origin | 64 | DS1
V5 | Inner wash valve | 65 | DS8
V27 | Vaccum tank waste valve | 65 | DS7
V7 | S Interior wash back valve | 65 | DS6
V4 | R2 Inner wash valve | 65 | DS5
V3 | R1 Inner wash valve | 65 | DS4
V30 | R2 Mix detergent discharge valve | 65 | DS3
V30 | R2 Mix wash water discharge valve | 65 | DS2
V2 | Detergen dilution valve | 65 | DS1
| Test counter | 66 | DS1
V28 | DI Water supply valve | 67 | DS8
P6 | Vacuum pump | 67 | DS5
P7 | Waste water pump | 67 | DS4
P3 | Degasser | 67 | DS1
V8 | S Outer wash valve 1 | 68 | DS8
V11 | S/R1 Mix detergent discharge valve | 68 | DS7
V6 | S/R1 Mix wash water dischage valve | 68 | DS6
V13 | Probe outer wash valve | 68 | DS5
V23 | Air release valve for mixture | 69 | DS8
V19 | Wash water dispense valve 4 | 69 | DS7
V18 | Wash water dispense valve 3 | 69 | DS6
V17 | Wash water dispense valve 2 | 69 | DS5
V16 | Wash water dispense valve 1 | 69 | DS4
V15 | Detergent dispense valve 2 | 69 | DS3
V14 | Detergent dispense valve 1 | 69 | DS2
V9 | S Outer wash valve 2 | 69 | DS1
V12 | S Outer wash tank valve | 6A | DS7
V22 | Wash water drain valve | 6A | DS6
V20 | Mixture drain valve | 6A | DS5
V21 | Overflow drain valve | 6A | DS4
V26 | Overflow aspiration valve | 6A | DS3
V25 | Air release valve for overflow | 6A | DS2
V24 | Mixture aspiration valve | 6A | DS1
P9 | Diluted detergent pump | 6B | DS3
P1 | Detergent dilution pump | 6B | DS2
P2 | DI Water supply pump | 6B | DS1
| Diluted detergent Dil. Det. EMPTY | 70 | DS8
| Diluted detergent Lower pos. | 70 | DS7
| Diluted detergent Upper pos. | 70 | DS6
| Diluted detergent Dil. Det. Overflow | 70 | DS5
| DI Water DI water Empty | 70 | DS4
| DI Water Lower pos. | 70 | DS3
| DI Water Upper pos. | 70 | DS2
| DI Water DI water Overflow | 70 | DS1
S4 | Vacuum tank full | 71 | DS5
S3 | Detergent tank Det. tank EMPTY | 71 | DS1
S6 | Conc. Waste water tank full | 72 | DS2
S5 | Washing waste water tank full | 72 | DS1
| Rear DIAG SW | 73 | DS5
| DIAG SW | 73 | DS2
| STAT SW | 73 | DS1
S8 | Weak negative-pressure | 75 | DS1

## AU27005400
AU2700/5400
No. | Short name | ADD | Ds
GAS01 | Cuv.Wheel Sensor Stop Pos | 43 | DS2
GAS00 | Cuv.Wheel Sensor Origin | 43 | DS1
FAS09 | Sample Probe Sensor Rotation BUSY | 46 | DS8
FAS19 | Sample Probe Sensor Up/Down BUSY | 46 | DS7
FAS01 | Sample Probe Sensor Stop Pos | 46 | DS6
FAS00 | Sample Probe Sensor Origin Pos | 46 | DS5
FAS02 | Sample Probe Sensor Rotation Pos | 46 | DS4
FAS21 | Sample Probe Sensor Crash | 46 | DS3
FAS12 | Sample Probe Sensor Level | 46 | DS2
FAS10 | Sample Probe Sensor Upper | 46 | DS1
FAS11 | Sample Probe Sensor Middle | 47 | DS1
FD11S09 | R11 Probe Sensor Rotation BUSY | 4B | DS8
FD11S19 | R11 Probe Sensor Up/Down BUSY | 4B | DS7
FD11S01 | R11 Probe Sensor Stop | 4B | DS6
FD11S00 | R11 Probe Sensor Origin Pos (Dsp.) | 4B | DS5
FD11S02 | R11 Probe Sensor Rotation Pos Pos (Asp.) | 4B | DS4
FD11S21 | R11 Probe Sensor Crash | 4B | DS3
FD11S12 | R11 Probe Sensor Level | 4B | DS2
FD11S10 | R11 Probe Sensor Upper | 4B | DS1
FD11S11 | R11 Probe Sensor Middle | 4C | DS1
FD01S09 | R12 Probe Sensor Rotation BUSY | 50 | DS8
FD01S19 | R12 Probe Sensor Up/Down BUSY | 50 | DS7
FD01S01 | R12 Probe Sensor Stop | 50 | DS6
FD01S00 | R12 Probe Sensor Origin Pos (Dsp.) | 50 | DS5
FD01S02 | R12 Probe Sensor Rotation Pos Pos (Asp.) | 50 | DS4
FD01S21 | R12 Probe Sensor Crash | 50 | DS3
FD01S12 | R12 Probe Sensor Level | 50 | DS2
FD01S10 | R12 Probe Sensor Upper | 50 | DS1
FD01S11 | R12 Probe Sensor Middle | 51 | DS1
FD21S09 | R21 Probe Sensor Rotation BUSY | 55 | DS8
FD21S19 | R21 Probe Sensor Up/Down BUSY | 55 | DS7
FD21S01 | R21 Probe Sensor Stop | 55 | DS6
FD21S00 | R21 Probe Sensor Origin Pos (Dsp.) | 55 | DS5
FD21S02 | R21 Probe Sensor Rotation Pos Pos (Asp.) | 55 | DS4
FD21S21 | R21 Probe Sensor Crash | 55 | DS3
FD21S12 | R21 Probe Sensor Level | 55 | DS2
FD21S10 | R21 Probe Sensor Upper | 55 | DS1
FD21S11 | R21 Probe Sensor Middle | 56 | DS1
FD31S09 | R22 Probe Sensor Rotation BUSY | 5A | DS8
FD31S19 | R22 Probe Sensor Up/Down BUSY | 5A | DS7
FD31S01 | R22 Probe Sensor Stop | 5A | DS6
FD31S00 | R22 Probe Sensor Origin Pos (Dsp.) | 5A | DS5
FD31S02 | R22 Probe Sensor Rotation Pos Pos (Asp.) | 5A | DS4
FD31S21 | R22 Probe Sensor Crash | 5A | DS3
FD31S12 | R22 Probe Sensor Level | 5A | DS2
FD31S10 | R22 Probe Sensor Upper | 5A | DS1
FD31S11 | R22 Probe Sensor Middle | 5B | DS1
DA01S09 | R1 Compartment Sensor Rotation BUSY | 5F | DS8
DA01S21 | R1 Compartment Sensor Bottle Detection | 5F | DS5
DA01S11 | R1 Compartment Sensor Cover | 5F | DS4
DA01S01 | R1 Compartment Sensor Stop | 5F | DS3
DA01S00 | R1 Compartment Sensor Origin | 5F | DS1
DA11S09 | R2 Compartment Sensor Rotation BUSY | 60 | DS8
DA11S21 | R2 Compartment Sensor Bottle Detection | 60 | DS5
DA11S11 | R2 Compartment Sensor Cover | 60 | DS4
DA11S01 | R2 Compartment Sensor Stop | 60 | DS3
DA11S00 | R2 Compartment Sensor Origin | 60 | DS1
SA21S01 | S1/R11/R21 Dispenser Sensor R21 Lower | 63 | DS6
SA21S00 | S1/R11/R21 Dispenser Sensor R21 Upper | 63 | DS5
SA01S01 | S1/R11/R21 Dispenser Sensor S1 Lower | 63 | DS4
SA01S01 | S1/R11/R21 Dispenser Sensor S1 Upper | 63 | DS3
SA11S01 | S1/R11/R21 Dispenser Sensor R11 Lower | 63 | DS2
SA11S00 | S1/R11/R21 Dispenser Sensor R11 Upper | 63 | DS1
SA21S09 | S1/R11/R21 Dispenser Sensor R21 BUSY | 64 | DS3
SA01S09 | S1/R11/R21 Dispenser Sensor S1 BUSY | 64 | DS2
SA11S09 | S1/R11/R21 Dispenser Sensor R11 BUSY | 64 | DS1
SA01S21 | S1/R11/R21 Dispenser Sensor Detection(Neg.) | 65 | DS2
SA01S22 | S1/R11/R21 Dispenser Sensor Detection(Pos.) | 65 | DS1
SA22S01 | S2/R12/R22 Dispenser Sensor R22 Lower | 68 | DS6
SA22S00 | S2/R12/R22 Dispenser Sensor R22 Upper | 68 | DS5
SA02S01 | S2/R12/R22 Dispenser Sensor S2 Lower | 68 | DS4
SA02S01 | S2/R12/R22 Dispenser Sensor S2 Upper | 68 | DS3
SA12S01 | S2/R12/R22 Dispenser Sensor R12 Lower | 68 | DS2
SA12S00 | S2/R12/R22 Dispenser Sensor R12 Upper | 68 | DS1
SA22S09 | S2/R12/R22 Dispenser Sensor R22 BUSY | 69 | DS3
SA02S09 | S2/R12/R22 Dispenser Sensor S2 BUSY | 69 | DS2
SA12S09 | S2/R12/R22 Dispenser Sensor R12 BUSY | 69 | DS1
SA02S21 | S2/R12/R22 Dispenser Sensor Detection(Neg.) | 6A | DS2
SA02S22 | S2/R12/R22 Dispenser Sensor Detection(Pos.) | 6A | DS1
FCS09 | MIX Sensor Rotation BUSY | 6E | DS8
FCS19 | MIX Sensor Up/Down BUSY | 6E | DS7
FCS01 | MIX Sensor Rotation Stop Pos | 6E | DS6
FCS00 | MIX Sensor Rotation Origin | 6E | DS5
FCS10 | MIX Sensor Upper | 6E | DS1
FCS29 | MIX Sensor Mixing BUSY | 6F | DS7
FBS09 | Cuv.Washer Sensor Up/Down BUSY | 75 | DS8
FAS02 | FA Rotation Pos 2 | 75 | DS7
FBS01 | Cuv.Washer Sensor Lower | 75 | DS2
FBS00 | Cuv.Washer Sensor Upper | 75 | DS1
CBS19 | Rack Supply Sensor M1 BUSY | 79 | DS8
CBS29 | Rack Supply Sensor M2 BUSY | 79 | DS7
CBS06 | Rack Supply Sensor Reverse Prohibition | 79 | DS4
CBS05 | Rack Supply Sensor Forward Prohibition | 79 | DS3
CBS01 | Rack Supply Sensor Feed | 79 | DS2
CBS02 | Rack Supply Sensor Receptacl | 79 | DS1
CES19 | Rack Retrieval Sensor Retrieval BUSY | 7A | DS8
CES29 | Rack Retrieval Sensor Carry-out BUSY | 7A | DS7
CES05 | Rack Retrieval Sensor Carry-Out LA Pos | 7A | DS6
CES04 | Rack Retrieval Sensor Carry-out Repeat Pos | 7A | DS5
CES03 | Rack Retrieval Sensor Carry-out Origin | 7A | DS4
CES02 | Rack Retrieval Sensor Retrieva Stop Pos | 7A | DS2
CES01 | Rack Retrieval Sensor Retrieva Origin | 7A | DS1
CFS19 | Rack Storage Sensor Retrieval BUSY | 7B | DS8
CAS08 | Rack Storage Sensor Rack Full | 7B | DS5
CFS02 | Rack Storage Sensor Stop Pos | 7B | DS2
CFS01 | Rack Storage Sensor Origin | 7B | DS1
CBS03 | Rack Type Detection Sensor Rack Present | 7C | DS5
CBS04-2 | Rack Type Detection Sensor Rack Type d2 | 7C | DS3
CBS04-1 | Rack Type Detection Sensor Rack Type d1 | 7C | DS2
CBS04-0 | Rack Type Detection Sensor Rack Type d0 | 7C | DS1
CAS07 | Rack Pass Detection Sensor CE Rack Pass 2 | 7D | DS4
CAS06 | Rack Pass Detection Sensor CE Rack Pass 1 | 7D | DS3
CAS05 | Rack Pass Detection Sensor CC Rack Pass 2 | 7D | DS2
CAS04 | Rack Pass Detection Sensor CC Rack Pass 1 | 7D | DS1
CBS07 | LA Feed-in/Carry-out Sensor Rack Reverse Prohibition | 7E | DS6
CBS10 | LA Feed-in/Carry-out Sensor Carry-out Detection | 7E | DS5
CBS09 | LA Feed-in/Carry-out Sensor Rack Pass | 7E | DS2
CBS08 | LA Feed-in/Carry-out Sensor Rack Detection | 7E | DS1
CCS19 | Rack Send-in Sensor Send-in BUSY | 80 | DS8
CCS29 | Rack Send-in Sensor Carry-out BUSY | 80 | DS7
CCS04 | Rack Send-in Sensor Carry-out Stop Pos | 80 | DS6
CCS03 | Rack Send-in Sensor Carry-out Origin | 80 | DS5
CAS03 | Rack Send-in Sensor Cup Detection | 80 | DS3
CCS02 | Rack Send-in Sensor Send-in Stop Pos | 80 | DS2
CCS01 | Rack Send-in Sensor Send-in Origin | 80 | DS1
CDS19 | Rack Buffer Sensor M1 BUSY | 81 | DS8
CDS29 | Rack Buffer Sensor M2 BUSY | 81 | DS7
CDS09 | Rack Buffer Sensor Reverse Prohibition | 81 | DS6
CDS08 | Rack Buffer Sensor Forward Prohibition | 81 | DS5
CDS02 | Rack Buffer Sensor Receptacle Pos | 81 | DS2
CDS01 | Rack Buffer Sensor Feed Pos | 81 | DS1
CLS19 | Rack Return Sensor Return BUSY | 82 | DS8
CLS02 | Rack Return Sensor Stop Pos | 82 | DS2
CLS01 | Rack Return Sensor Origin Pos | 82 | DS1
CDS04 | Rack Type Detection Sensor Rack Present 2 | 83 | DS6
CDS03 | Rack Type Detection Sensor Rack Present 1 | 83 | DS5
CAS10-2 | Rack Type Detection Sensor Rack Type-2 | 83 | DS3
CAS10-1 | Rack Type Detection Sensor Rack Type-1 | 83 | DS2
CAS10-1 | Rack Type Detection Sensor Rack Type-0 | 83 | DS1
CDS07 | Rack Pass Detection Sensor ANL Rack3 | 84 | DS3
CDS06 | Rack Pass Detection Sensor ANL Rack2 | 84 | DS2
CDS05 | Rack Pass Detection Sensor ANL Rack1 | 84 | DS1
DCS09 | STAT Sensor Rotaton BUSY | 87 | DS8
DCS11 | STAT Sensor Cover | 87 | DS5
DCS21 | STAT Sensor Cup Detection | 87 | DS3
DCS01 | STAT Sensor Stop Pos | 87 | DS2
DCS00 | STAT Sensor Origin | 87 | DS1
CGS19 | Rack Transfer1 Sensor Belt BUSY | 88 | DS8
CGS29 | Rack Transfer1 Sensor Lever BUSY | 88 | DS7
CGS03 | Rack Transfer1 Sensor Rack Step Detection 1 | 88 | DS3
CGS02 | Rack Transfer1 Sensor Lever 1 Stop Pos | 88 | DS2
CGS01 | Rack Transfer1 Sensor Lever 1 Origin | 88 | DS1
CGS39 | Rack Transfer1 Sensor Belt BUSY | 89 | DS8
CGS49 | Rack Transfer1 Sensor Lever BUSY | 89 | DS7
CGS07 | Rack Transfer1 Sensor Rack Step Detection 2 | 89 | DS3
CGS06 | Rack Transfer1 Sensor Lever 2 Stop Pos | 89 | DS2
CGS05 | Rack Transfer1 Sensor Lever 2 Origin | 89 | DS1
DCV35 | Degassed water supply valve | 0A | DS3
DCV34 | Detergent B dilution valve | 0A | DS2
DCV33 | Deionized water bubble without | 0A | DS1
DCP08 | Pure detergent A supply pump | 0C | DS5
DCV37 | Inner wash water return valve 2 | 13 | DS8
DCV07 | Inner wash water return valve | 13 | DS7
DCV06 | R22 inner wsah valve | 13 | DS6
DCV05 | R21 inner wsah valve | 13 | DS5
DCV04 | R12 inner wsah valve | 13 | DS4
DCV03 | R11 inner wsah valve | 13 | DS3
DCV02 | S2 inner wsah valve | 13 | DS2
DCV01 | S1 inner wsah valve | 13 | DS1
DCV15 | Liquid dripping drain valve | 14 | DS8
DCV14 | Mixture drain valve | 14 | DS7
DCV13 | Overflow drain valve | 14 | DS6
DCV12 | Wash water drain valve | 14 | DS5
DCV32 | Detergent A dilution valve | 14 | DS3
DCV09 | S2/R12/R2 outer wash valve | 14 | DS2
DCV08 | S1/R11/R12 outer wash valve | 14 | DS1
DCV30 | Detergent A dispense valve | 15 | DS8
DCV29 | Detergent B dispense valve | 15 | DS7
DCV28 | Wash water dispense valve2 | 15 | DS6
DCV27 | Wash water dispense valve1 | 15 | DS5
DCV26 | Detergent A drip prevention valve | 15 | DS3
DCV25 | Detergent B drip prevention valve | 15 | DS2
DCV24 | Wash water drip prevention valve | 15 | DS1
DCV23 | Drip air release valve | 16 | DS8
DCV22 | Mixture air release valve | 16 | DS7
DCV21 | Overflow air release valve | 16 | DS6
DCV20 | Aspiration air release valve | 16 | DS5
DCV19 | Liquid dripping aspiration valve | 16 | DS4
DCV18 | Mixture aspiration valve | 16 | DS3
DCV17 | Overflow aspiration valve | 16 | DS2
DCV16 | Wash water aspiration valve | 16 | DS1
DCP09 | DI water supply pump | 17 | DS7
DCP10 | Detergent B supply pump | 17 | DS6
DCP07 | Detergent A supply pump | 17 | DS5
DCV11 | Mix detergent valve | 17 | DS2
DCV10 | Mix wash valve | 17 | DS1
DCP11 | Pure detergent B pump | 18 | DS8
DCP12 | Degassed water rotation pump | 18 | DS7
ACV31 | Vacuum tank waste valve | 18 | DS6
ACP13 | Waste pump | 18 | DS5
DEGASSOR | Degasser | 18 | DS4
ACP06 | Vacuum pump 2 | 18 | DS2
ACP05 | Vacuum pump 1 | 18 | DS1

## AU5400
AU5400 ISE
No. | Short name | ADD | Ds
DCV03 | Washer pump | 0C | DS3
DCV02 | S outer wsah valve | 0C | DS2
DCV01 | S inner wash valve | 0C | DS1
FAS09 | Sample Probe Sensor Rotation BUSY | 42 | DS8
FAS19 | Sample Probe Sensor Up/Down BUSY | 42 | DS7
FAS01 | Sample Probe Sensor Stop Pos | 42 | DS6
FAS00 | Sample Probe Sensor Origin Pos | 42 | DS5
FAS02 | Sample Probe Sensor Rotation Pos | 42 | DS4
FAS21 | Sample Probe Sensor Crash | 42 | DS3
FAS21 | Sample Probe Sensor Level | 42 | DS2
FAS10 | Sample Probe Sensor Upper | 42 | DS1
FAS11 | Sample Probe Sensor M iddle | 43 | DS1
SA21S01 | ISE Dispenser Sensor R21 Lower | 46 | DS6
SA21S00 | ISE Dispenser Sensor R21 U pper | 46 | DS5
SA01S01 | ISE Dispenser Sensor S1 Lower | 46 | DS4
SA01S00 | ISE Dispenser Sensor S1 U pper | 46 | DS3
SA11S01 | ISE Dispenser Sensor R11 Lower | 46 | DS2
SA11S00 | ISE Dispenser Sensor R11 U pper | 46 | DS1
SA21S09 | ISE Dispenser Sensor R21 BUSY | 47 | DS3
SA01S09 | ISE Dispenser Sensor S1 BUSY | 47 | DS2
SA11S09 | ISE Dispenser Sensor R11 BUSY | 47 | DS1
SA01S21 | ISE Dispenser C log Sensor Detection(N eg.) | 48 | DS2
SA01S22 | ISE Dispenser C log Sensor Detection(P os.) | 48 | DS1
AU5400 SMP
No. | Short name | ADD | Ds
DCV35 | Degassed water supply valve | 0A | DS3
DCV34 | Detergent B dilution valve | 0A | DS2
DCV33 | Deionized water bubble without | 0A | DS1
CBS06 | Rack Supply Sensor Reverse Prohibition | 41 | DS4
CBS05 | Rack Supply Sensor Forward Prohibition | 41 | DS3
CBS01 | Rack Supply Sensor Feed | 41 | DS2
CBS02 | Rack Supply Sensor Receptacl | 41 | DS1
CES19 | Rack Retrieval Sensor Retrieval BUSY | 42 | DS8
CES29 | Rack Retrieval Sensor Carry-out BUSY | 42 | DS7
CES05 | Rack Retrieval Sensor Carry-Out LA Pos | 42 | DS6
CES04 | Rack Retrieval Sensor Carry-out Repeat Pos | 42 | DS5
CES03 | Rack Retrieval Sensor Carry-Out Origin | 42 | DS4
CES02 | Rack Retrieval Sensor Retrieval Stop Pos | 42 | DS2
CES01 | Rack Retrieval Sensor Retrieval Origin | 42 | DS1
CAS08 | Rack Storage Sensor Rack Full | 43 | DS5
CFS02 | Rack Storage Sensor Stop Pos | 43 | DS2
CFS01 | Rack Storage Sensor Origin | 43 | DS1
CBS03 | Rack Type Detection Sensor Rack Present | 44 | DS5
CBS04-2 | Rack Type Detection Sensor Rack Type-2 | 44 | DS3
CBS04-1 | Rack Type Detection Sensor Rack Type-1 | 44 | DS2
CBS04-0 | Rack Type Detection Sensor Rack Type-0 | 44 | DS1
CAS07 | Rack Pass Detection Sensor CE Rack Pass 2 | 45 | DS4
CAS06 | Rack Pass Detection Sensor CE Rack Pass 1 | 45 | DS3
CAS05 | Rack Pass Detection Sensor CC Rack Pass 2 | 45 | DS2
CAS04 | Rack Pass Detection Sensor CE Rack Pass 1 | 45 | DS1
CBS07 | LA Feed-in/Carry-out Sensor Rack Reverse Prohibition | 46 | DS6
CBS10 | LA Feed-in/Carry-out Sensor Carry-out Detection | 46 | DS5
CBS09 | LA Feed-in/Carry-out Sensor Rack Pass | 46 | DS2
CBS08 | LA Feed-in/Carry-out Sensor Rack Detection | 46 | DS1
CCS04 | Rack Send-in Sensor Carry-out Stop Pos | 48 | DS6
CCS03 | Rack Send-in Sensor Carry-out Origin | 48 | DS5
CAS03 | Rack Send-in Sensor Cup Detection | 48 | DS3
CCS02 | Rack Send-in Sensor Send-in Stop Pos | 48 | DS2
CCS01 | Rack Send-in Sensor Send-in Origin | 48 | DS1
CDS09 | Rack Buffer Sensor Reverse Prohibition | 49 | DS6
CDS08 | Rack Buffer Sensor Forward Prohibition | 49 | DS5
CDS02 | Rack Buffer Sensor Feed | 49 | DS2
CDS01 | Rack Buffer Sensor Receptacle | 49 | DS1
CBS11 | Rack Return Sensor LA Manual Rack | 4A | DS3
CLS02 | Rack Return Sensor Stop Pos | 4A | DS2
CLS01 | Rack Return Sensor Origin Pos | 4A | DS1
CDS04 | Rack Type Detection Sensor Rack Present 2 | 4B | DS6
CDS03 | Rack Type Detection Sensor Rack Present 1 | 4B | DS5
CAS10-2 | Rack Type Detection Sensor Rack Type-2 | 4B | DS3
CAS10-1 | Rack Type Detection Sensor Rack Type-1 | 4B | DS2
CAS10-0 | Rack Type Detection Sensor Rack Type-0 | 4B | DS1
CDS07 | Rack Pass Detection Sensor ANL Rack Pass3 | 4C | DS3
CDS06 | Rack Pass Detection Sensor ANL Rack Pass2 | 4C | DS2
CDS05 | Rack Pass Detection Sensor ANL Rack Pass1 | 4C | DS1
CHS14 | ANL1 Retrieval Lane (CH3) Lane Rack Pass | 4E | DS6
CHS15 | ANL1 Retrieval Lane (CH3) Rack Present | 4E | DS5
CHS04 | ANL1 Passing Lane (CH1) Lane Rack Pass | 4F | DS6
CHS02 | ANL1 Passing Lane (CH1) Stop Pos | 4F | DS2
CHS01 | ANL1 Passing Lane (CH1) Origin | 4F | DS1
CHS13 | ANL1 Dispensing Lane(CH2) Rack present After Dispense | 50 | DS6
CHS07 | ANL1 Dispensing Lane(CH2) Rack present Pre .Dispence | 50 | DS5
CHS09 | ANL1 Dispensing Lane(CH2) Stop Pos | 50 | DS2
CHS08 | ANL1 Dispensing Lane(CH2) Origin | 50 | DS1
CHS17 | ANL1 Lever Open/Close Stop Pos | 51 | DS6
CHS16 | ANL1 Lever Open/Close Origin | 51 | DS5
CHS03 | ANL1 Lever Open/Close Lane 1 Rack Step | 51 | DS2
CHS10 | ANL1 Lever Open/Close Lane 2 Rack Step | 51 | DS1
CJS07 | ANL1 Retrieval Lane(CJ3) Left side Rack Pass | 53 | DS3
CJFS06 | ANL1 Retrieval Lane(CJ3) Right side Rack Pass | 53 | DS1
CJS02 | ANL1 Retrieval Lane(CJ1) Left side Rack Pass | 54 | DS3
CJFS01 | ANL1 Retrieval Lane(CJ1) Right side Rack Pass | 54 | DS1
CJS05 | ANL1 Retrieval Lane(CJ2) Waiting Pos.2 Rack | 55 | DS3
CJS04 | ANL1 Retrieval Lane(CJ2) Waiting Pos.3 Rack | 55 | DS2
CJFS03 | ANL1 Retrieval Lane(CJ2) Right side Rack Pass | 55 | DS1
CJS09 | ANL1 Lane change.Stopper CJ Stopper Stop Pos. | 56 | DS6
CJS08 | ANL1 Lane change.Stopper CJ Stopper Origin | 56 | DS5
CJS04 | ANL1 Lane change.Stopper Retrieval Lane Rack | 56 | DS4
CJS03 | ANL1 Lane change.Stopper CK Retrieval Lane Pos. | 56 | DS3
CJS02 | ANL1 Lane change.Stopper CK Dispensing Lane Pos. | 56 | DS2
CJS01 | ANL1 Lane change.Stopper CK Origin | 56 | DS1
CHS12 | ANL1 Stopper Stopper Stop Pos.After Dispense | 58 | DS6
CHS11 | ANL1 Stopper Stopper Origin After Dispense | 58 | DS5
CHS06 | ANL1 Stopper Stopper Stop Pos.Pre. Dispense | 58 | DS2
CHS05 | ANL1 Stopper Stopper Origin Pre. Dispense | 58 | DS1
CHS14 | ANL2 Retrieval Lane(CH3) Lane Rack Pass | 5A | DS6
CHS15 | ANL2 Retrieval Lane(CH3) Rack Present | 5A | DS5
CHS04 | ANL2 Passing Lane(CH1) Lane1 Rack PASS | 5B | DS6
CHS02 | ANL2 Passing Lane(CH1) Stop Pos | 5B | DS2
CHS01 | ANL2 Passing Lane(CH1) Origin | 5B | DS1
CHS13 | ANL2 Dispensing Lane(CH2) Rack Present After Dispense | 5C | DS6
CHS07 | ANL2 Dispensing Lane(CH2) Rack Present Pre.Dispence | 5C | DS5
CHS09 | ANL2 Dispensing Lane(CH2) Stop Pos | 5C | DS2
CHS08 | ANL2 Dispensing Lane(CH2) Origin | 5C | DS1
CHS17 | ANL2 Lever Open/Close Stop Pos | 5D | DS6
CHS16 | ANL2 Lever Open/Close Origin | 5D | DS5
CHS03 | ANL2 Lever Open/Close Lane1 Rack STEP | 5D | DS2
CHS10 | ANL2 Lever Open/Close Lane2 Rack STEP | 5D | DS1
CJS07 | ANL2 Retrieval Lane(CJ3) Left side Rack Pass | 5F | DS3
CJFS06 | ANL2 Retrieval Lane(CJ3) Right side Rack Pass | 5F | DS1
CJS02 | ANL2 Passing Lane(CH1) Left side Rack Pass | 60 | DS3
CJFS01 | ANL2 Passing Lane(CH1) Right side Rack Pass | 60 | DS1
CJS05 | ANL2 Dispensing Lane(CJ2) Waiting Pos.2 RACK | 61 | DS3
CJS04 | ANL2 Dispensing Lane(CJ2) Waiting Pos.3 RACK | 61 | DS2
CJFS03 | ANL2 Dispensing Lane(CJ2) Right side Rack Pass | 61 | DS1
CJS09 | ANL2 Lane change.Stopper CJ Stopper Stop Pos | 62 | DS6
CJS08 | ANL2 Lane change.Stopper CJ Stopper Origin | 62 | DS5
CKS04 | ANL2 Lane change.Stopper Retrieval Lane RACK | 62 | DS4
CKS03 | ANL2 Lane change.Stopper CK Retrieval Lane Pos | 62 | DS3
CKS02 | ANL2 Lane change.Stopper CK Dispensing Lane Pos | 62 | DS2
CKS01 | ANL2 Lane change.Stopper CK Origin | 62 | DS1
CHS12 | ANL2 Stopper Stopper Stop Pos.After Dispense | 64 | DS6
CHS11 | ANL2 Stopper Stopper Origin After Dispense | 64 | DS5
CHS06 | ANL2 Stopper Stopper Stop Pos.Pre.Dispense | 64 | DS2
CHS05 | ANL2 Stopper Stopper Origin Pre.Dispense | 64 | DS1
CHS14 | ANL3 Retrieval Lane(CH3) Lane Rack PASS | 66 | DS6
CHS15 | ANL3 Retrieval Lane(CH3) Rack Present | 66 | DS5
CHS04 | ANL3 Passing Lane(CH1) Lane1 Rack PASS | 67 | DS6
CHS02 | ANL3 Passing Lane(CH1) Stop Pos. | 67 | DS2
CHS01 | ANL3 Passing Lane(CH1) Origin | 67 | DS1
CHS13 | ANL3 Dispensing Lane(CH2) Rack Present After Dispense | 68 | DS6
CHS07 | ANL3 Dispensing Lane(CH2) Rack Present Pre.Dispence | 68 | DS5
CHS09 | ANL3 Dispensing Lane(CH2) Stop Pos. | 68 | DS2
CHS08 | ANL3 Dispensing Lane(CH2) Origin | 68 | DS1
CHS17 | ANL2 Lever Open/Close Stop Pos. | 69 | DS6
CHS16 | ANL2 Lever Open/Close Origin | 69 | DS5
CHS03 | ANL2 Lever Open/Close Lane1 Rack STEP | 69 | DS2
CHS10 | ANL2 Lever Open/Close Lane2 Rack STEP | 69 | DS1
CJS07 | ANL3 Retrieval Lane(CJ3) Left side Rack Pass | 6B | DS3
CJFS06 | ANL3 Retrieval Lane(CJ3) Right side Rack Pass | 6B | DS1
CJS02 | ANL3 Passing Lane(CJ) Left side Rack Pass | 6C | DS3
CJFS01 | ANL3 Passing Lane(CJ) Right side Rack Pass | 6C | DS1
CJS05 | ANL3 Dispensing Lane(CJ2) Waiting Pos.2 RACK | 6D | DS3
CJS04 | ANL3 Dispensing Lane(CJ2) Waiting Pos.3 RACK | 6D | DS2
CJFS03 | ANL3 Dispensing Lane(CJ2) Right side Rack Pass | 6D | DS1
CJS09 | ANL3 Lane change.Stopper CJ Stopper Stop Pos | 6E | DS6
CJS08 | ANL3 Lane change.Stopper CJ Stopper Origin | 6E | DS5
CKS04 | ANL3 Lane change.Stopper Retrieval Lane RACK | 6E | DS4
CKS03 | ANL3 Lane change.Stopper CK Retrieval Lane Pos | 6E | DS3
CKS02 | ANL3 Lane change.Stopper CK Dispensing Lane Pos | 6E | DS2
CKS01 | ANL3 Lane change.Stopper CK Origin | 6E | DS1
CHS12 | ANL3 Stopper Stopper Stop Pos.After Dispense | 70 | DS6
CHS11 | ANL3 Stopper Stopper Origin After Dispense | 70 | DS5
CHS06 | ANL3 Stopper Stopper Stop Pos.Pre.Dispense | 70 | DS2
CHS05 | ANL3 Stopper Stopper Origin Pre.Dispense | 70 | DS1
CMS09 | ISE R etrieval Lane (CM3) Lane3 Rack Pass | 72 | DS6
CMS10 | ISE R etrieval Lane (CM3) Rack Present | 72 | DS5
CMS04 | ISE D ispens ing Lane (CM1) Lane1 Rack PASS | 73 | DS6
CMS02 | ISE D ispens ing Lane (CM1) Stop Pos | 73 | DS2
CMS01 | ISE D ispens ing Lane (CM1) Origin | 73 | DS1
CMS08 | ISE D ispens ing Lane (CM2) Rack present After Dispense | 74 | DS6
CMS06 | ISE D ispens ing Lane (CM2) Stop Pos | 74 | DS2
CMS05 | ISE D ispens ing Lane (CM2) Origin | 74 | DS1
CNS01 | Rack receive,Rack vertical feed Arrival detection | 7C | DS5
CNS10 | Rack receive,Rack vertical feed Backward prohibition | 7C | DS4
CNS11 | Rack receive,Rack vertical feed Forward prohibition | 7C | DS3
CNS09 | Rack receive,Rack vertical feed Feed pos | 7C | DS2
CNS08 | Rack receive,Rack vertical feed Receive pos | 7C | DS1
CNS07 | Rack push Rack pass | 7F | DS4
CNS04 | Rack push Rack detection | 7F | DS3
CNS03 | Rack push Stop pos | 7F | DS2
CNS02 | Rack push Origin | 7F | DS1
CNS06 | Horizontal feed Rack Stop Pos | 80 | DS2
CNS05 | Horizontal feed Rack Origin | 80 | DS1
CNS12 | Rack carryout Carry out Rack pass | 81 | DS3
CNS14 | Rack carryout Stop pos | 81 | DS2
CNS13 | Rack carryout Origin | 81 | DS1
CMS03 | ISE Lane Lane1 Rack Step | 84 | DS2
CMS07 | ISE Lane Lane2 Rack Step | 84 | DS1

## Sheet7
AU5800 ANL
No. | Short name | ADD | Ds
DAS02 | R1 REF (DA01) / Rotation Stop sensor | 02 | DS7
DAS04 | R1 REF (DA01) / Rotation Bottle sensor | 02 | DS3
DAS03 | R1 REF (DA01) / Rotation Cover sensor | 02 | DS2
DAS01 | R1 REF (DA01) / Rotation Origin sensor | 02 | DS1
DAS12 | R2 REF (DA11) / Rotation Stop sensor | 06 | DS7
DAS14 | R2REF (DA11) / Rotation Bottle sensor | 06 | DS3
DAS13 | R2 REF (DA11) / Rotation Cover sensor | 06 | DS2
DAS11 | R2 REF (DA11) / Rotation Origin sensor | 06 | DS1
FAS10 | INNER S transfer unit (FA01)/ Up down S probe Liquid detection sensor | 0A | DS5
FAS09 | INNER S transfer unit (FA01)/ Up down S probe crash detection sensor | 0A | DS4
FAS01 | INNER S transfer unit (FA01)/ Up down Upper sensor | 0A | DS3
FAS02 | INNER S transfer unit (FA01)/ Up down Middle 1 sensor | 0A | DS1
FAS08 | INNER S transfer unit (FA01)/ Rotation Rotation area 2 sensor | 0E | DS5
FAS07 | INNER S transfer unit (FA01)/ Rotation Rotation area 1 sensor | 0E | DS4
FAS06 | INNER S transfer unit (FA01)/ Rotation Rotation stop 2 sensor | 0E | DS3
FAS05 | INNER S transfer unit (FA01)/ Rotation Rotation stop 1 sensor | 0E | DS2
FAS04 | INNER S transfer unit (FA01)/ Rotation Rotation origin sensor | 0E | DS1
FAS20 | OUTER S transfer unit (FA11)/ Up down S probe Liquid detection sensor | 10 | DS5
FAS19 | OUTER S transfer unit (FA11)/ Up down S probe crash detection sensor | 10 | DS4
FAS11 | OUTER S transfer unit (FA11)/ Up down Upper sensor | 10 | DS3
FAS12 | OUTER S transfer unit (FA11)/ Up down Middle 1 sensor | 10 | DS1
FAS18 | OUTER S transfer unit (FA11)/ Rotation Rotation area 2 sensor | 14 | DS5
FAS17 | OUTER S transfer unit (FA11)/ Rotation Rotation area 1 sensor | 14 | DS4
FAS16 | OUTER S transfer unit (FA11)/ Rotation Rotation stop 2 sensor | 14 | DS3
FAS15 | OUTER S transfer unit (FA11)/ Rotation Rotation stop 1 sensor | 14 | DS2
FAS14 | OUTER S transfer unit (FA11)/ Rotation Rotation origin sensor | 14 | DS1
CJS03 | Sample Shutter unit(CJ01) Stop sensor 2 (Passing lane) | 16 | DS3
CJS02 | Sample Shutter unit(CJ01) Stop sensor 1(Dispensing lane) | 16 | DS2
CJS01 | Sample Shutter unit(CJ01) Origin sensor (Refuge position) | 16 | DS1
FBS02 | Cuvette Wash unit(FB01) / Up down Lower sensor | 18 | DS2
FBS01 | Cuvette Wash unit(FB01) / Up down Upper sensor | 18 | DS1
FCS02 | S/R Mix unit (FC01)/ Up down Lower sensor | 1B | DS2
FCS01 | S/R Mix unit (FC01)/ Up down Upper sensor | 1B | DS1
FCS04 | S/R Mix unit (FC01)/ Rotation Rotation stop sensor | 1E | DS7
FCS03 | S/R Mix unit (FC01)/ Rotation Rotation origin sensor | 1E | DS1
FCS05 | S/R Mix unit (FC01)/ Rotation Monitor sensor | 21 | DS7
FCS12 | R2 Mix unit (FC11)/ Up down | 24 | DS2
FCS11 | R2 Mix unit (FC11)/ Up down | 24 | DS1
FCS14 | R2 Mix unit (FC11)/ Rotation | 27 | DS7
FCS13 | R2 Mix unit (FC11)/ Rotation | 27 | DS1
FCS15 | R2 Mix unit (FC11)/ Rotation Monitor sensor | 2A | DS7
FDS10 | INNER R1 Transfer unit (FD01)/ Up down R1 Probe liquid detection sensor | 2D | DS5
FDS09 | INNER R1 Transfer unit (FD01)/ Up down R1 Probe crash detection sensor | 2D | DS4
FDS01 | INNER R1 Transfer unit (FD01)/ Up down Upper sensor | 2D | DS3
FDS03 | INNER R1 Transfer unit (FD01)/ Up down Middle 2 sensor | 2D | DS2
FDS02 | INNER R1 Transfer unit (FD01)/ Up down Middle 1 sensor | 2D | DS1
FDS08 | INNER R1 Transfer unit (FD01)/ Up down Rotation area 2 sensor | 31 | DS5
FDS07 | INNER R1 Transfer unit (FD01)/ Up down Rotation area 1 sensor | 31 | DS4
FDS06 | INNER R1 Transfer unit (FD01)/ Up down Rotation stop 2 sensor | 31 | DS3
FDS05 | INNER R1 Transfer unit (FD01)/ Up down Rotation stop 1 sensor | 31 | DS2
FDS04 | INNER R1 Transfer unit (FD01)/ Up down Rotation origin sensor | 31 | DS1
FDS20 | OUTER R1 Transfer unit (FD11)/ Up down R1 Probe liquid detection sensor | 34 | DS5
FDS19 | OUTER R1 Transfer unit (FD11)/ Up down R1 Probe crash detection sensor | 34 | DS4
FDS11 | OUTER R1 Transfer unit (FD11)/ Up down Upper sensor | 34 | DS3
FDS13 | OUTER R1 Transfer unit (FD11)/ Up down Middle 2 sensor | 34 | DS2
FDS12 | OUTER R1 Transfer unit (FD11)/ Up down Middle 1 sensor | 34 | DS1
FDS18 | OUTER R1 Transfer unit (FD11)/ Up down Rotation area 2 sensor | 37 | DS5
FDS17 | OUTER R1 Transfer unit (FD11)/ Up down Rotation area 1 sensor | 37 | DS4
FDS16 | OUTER R1 Transfer unit (FD11)/ Up down Rotation stop 2 sensor | 37 | DS3
FDS15 | OUTER R1 Transfer unit (FD11)/ Up down Rotation stop 1 sensor | 37 | DS2
FDS14 | OUTER R1 Transfer unit (FD11)/ Up down Rotation origin sensor | 37 | DS1
FDS30 | INNER R2Transfer unit (FD21)/ Up down R1 Probe liquid detection sensor | 39 | DS5
FDS29 | INNER R2 Transfer unit (FD21)/ Up down R1 Probe crash detection sensor | 39 | DS4
FDS21 | INNER R2 Transfer unit (FD21)/ Up down Upper sensor | 39 | DS3
FDS23 | INNER R2 Transfer unit (FD21)/ Up down Middle 2 sensor | 39 | DS2
FDS22 | INNER R2 Transfer unit (FD21)/ Up down Middle 1 sensor | 39 | DS1
FDS28 | INNER R2 Transfer unit (FD21)/ Up down Rotation area 2 sensor | 3D | DS5
FDS27 | INNER R2 Transfer unit (FD21)/ Up down Rotation area 1 sensor | 3D | DS4
FDS26 | INNER R2 Transfer unit (FD21)/ Up down Rotation stop 2 sensor | 3D | DS3
FDS25 | INNER R2 Transfer unit (FD21)/ Up down Rotation stop 1 sensor | 3D | DS2
FDS24 | INNER R2 Transfer unit (FD21)/ Up down Rotation origin sensor | 3D | DS1
FDS40 | OUTER R2 Transfer unit (FD31)/ Up down R1 Probe liquid detection sensor | 3F | DS5
FDS39 | OUTER R2 Transfer unit (FD31)/ Up down R1 Probe crash detection sensor | 3F | DS4
FDS31 | OUTER R2 Transfer unit (FD31)/ Up down Upper sensor | 3F | DS3
FDS33 | OUTER R2 Transfer unit (FD31)/ Up down Middle 2 sensor | 3F | DS2
FDS32 | OUTER R2 Transfer unit (FD31)/ Up down Middle 1 sensor | 3F | DS1
FDS38 | OUTER R2 Transfer unit (FD31)/ Up down Rotation area 2 sensor | 43 | DS5
FDS37 | OUTER R2 Transfer unit (FD31)/ Up down Rotation area 1 sensor | 43 | DS4
FDS36 | OUTER R2 Transfer unit (FD31)/ Up down Rotation stop 2 sensor | 43 | DS3
FDS35 | OUTER R2 Transfer unit (FD31)/ Up down Rotation stop 1 sensor | 43 | DS2
FDS34 | OUTER R2 Transfer unit (FD31)/ Up down Rotation origin sensor | 43 | DS1
SAS02 | INNER S Dispenser(SA02)Lower sensor | 45 | DS2
SAS01 | INNER S Dispenser(SA02)Upper sensor | 45 | DS1
SAS06 | INNER R1 Dispenser(SA01)Lower sensor | 47 | DS2
SAS05 | INNER R1 Dispenser(SA01)Upper sensor | 47 | DS1
SAS10 | INNER R2 Dispenser(SA01)Lower sensor | 49 | DS2
SAS09 | INNER R2 Dispenser(SA01)Upper sensor | 49 | DS1
SAS04 | OUTER S Dispenser(SA02)Lower sensor | 4B | DS2
SAS03 | OUTER S Dispenser(SA02)Upper sensor | 4B | DS1
SAS08 | OUTER R1 Dispenser(SA01)Lower sensor | 4F | DS2
SAS07 | OUTER R1 Dispenser(SA01)Upper sensor | 4F | DS1
SAS12 | OUTER R2 Dispenser(SA01)Lower sensor | 51 | DS2
SAS11 | OUTER R2 Dispenser(SA01)Upper sensor | 51 | DS1
SAS14 | INNER inside wash Dispenser (SA01)Lower sensor | 53 | DS2
SAS13 | INNER inside wash Dispenser (SA01)Upper sensor | 53 | DS1
SAS16 | OUTER inside wash Dispenser (SA01)Lower sensor | 55 | DS2
SAS15 | OUTER inside wash Dispenser (SA01)Upper sensor | 55 | DS1
GAS02 | Cuvette Wheel(GA01) / Rotation Stop sensor | 5D | DS2
GAS01 | Cuvette Wheel(GA01) / Rotation Origin sensor | 5D | DS1
AAV31 | ANL Main body (AA01)Mixture aspiration valve (V31) | 5F | DS8
AAV30 | ANL Main body (AA01)Air release valve for mixture (V30) | 5F | DS7
AAV29 | ANL Main body (AA01) Overflow aspiration valve(V29) | 5F | DS6
AAV28 | ANL Main body (AA01)Air release valve for overflow(V28) | 5F | DS5
AAV32 | ANL Main body (AA01)Vacuum tank drain valve(V32) | 5F | DS4
AAV25 | ANL Main body (AA01)Overflow drain valve (V25) | 5F | DS3
AAV02 | ANL Main body (AA01)OUTER S probe inside wash valve (V2) | 5F | DS2
AAV01 | ANL Main body (AA01)INNER S probe inside wash valve (V1) | 5F | DS1
AAV06 | ANL Main body (AA01)OUTER R2 probe inside wash valve (V6) | 60 | DS8
AAV05 | ANL Main body (AA01)INNER R2 probe inside wash valve (V5) | 60 | DS7
AAV04 | ANL Main body (AA01)OUTER R1 probe inside wash valve (V4) | 60 | DS6
AAV03 | ANL Main body (AA01)INNER R1 probe inside wash valve (V3) | 60 | DS5
AAV40 | ANL Main body (AA01)OUTER R probe inside switch valve (V40) | 60 | DS4
AAV39 | ANL Main body (AA01)INNER R probe inside switch valve (V39) | 60 | DS3
AAV38 | ANL Main body (AA01)S/R probe inside wash return valve (V38) | 60 | DS2
AAV38 | ANL Main body (AA01)Detergent dilution valve (V34) | 60 | DS1
(V28) | ANL Main body (AA01) D.I. water supply valve 1 CAV01 | 61 | DS8
OPDV01 | ANL Main body (AA01)D.I water supply option valve | 61 | DS7
AAP12 | ANL Main body (AA01)Vacuum pump 2 (P12) | 61 | DS6
AAP11 | ANL Main body (AA01)Vacuum pump 1(P11) | 61 | DS5
AAP08 | ANL Main body (AA01)Waste water pump (P8) | 61 | DS4
AAP14 | ANL Main body (AA01)Degasser pump(P14) | 61 | DS2
AAV33 | ANL Main body (AA01)D.I water supply valve 2 (V33) | 61 | DS1
AAV18 | ANL Main body (AA01)OUTER cuvette wash detergent valve 1 (V18) | 62 | DS8
AAV15 | ANL Main body (AA01)INNER cuvette wash water valve 3 (V15) | 62 | DS7
AAV13 | ANL Main body (AA01)INNER cuvette wash water valve 1 (V13) | 62 | DS6
AAV13 | ANL Main body (AA01)INNER cuvette wash detergent valve 1 (V11) | 62 | DS5
AAV26 | ANL Main body (AA01)Wash water drain valve V26) | 63 | DS8
AAV37 | ANL Main body (AA01) S probe outside wash valve 2 for HbA1c (V37) | 63 | DS7
AAV36 | ANL Main body (AA01)S probe outside wash peservoir valve for HbA1c (V36) | 63 | DS6
AAV35 | ANL Main body (AA01)S probe outside wash valve 1 for HbA1c (V35) | 63 | DS5
AAV24 | ANL Main body (AA01)OUTER cuvette wash water valve 5 (V24) | 63 | DS4
AAV17 | ANL Main body (AA01)INNER cuvette wash water valve 5 (V17) | 63 | DS3
AAV22 | ANL Main body (AA01)OUTER cuvette wash water valve 3 (V22) | 63 | DS2
AAV20 | ANL Main body (AA01)OUTER cuvette wash water valve 1 (V20) | 63 | DS1
AAV41 | ANL Main body (AA01)R1 probe cold water valve (V41) | 64 | DS6
AAV10 | ANL Main body (AA01)Mixing bar detergent valve (V10) | 64 | DS5
AAV09 | ANL Main body (AA01)Mixing bar wash water valve (V9) | 64 | DS4
AAV08 | ANL Main body (AA01)OUTER S/R probe outside wash valve (V8) | 64 | DS3
AAV07 | ANL Main body (AA01)INNER S/R probe outside wash valve (V7) | 64 | DS2
AAV27 | ANL Main body (AA01)Mixture drain valve (V27) | 64 | DS1
AAP07 | ANL Main body (AA01)Replenishment detergent pump (P7) | 65 | DS6
AAP06 | ANL Main body (AA01)Detergent pump(P6) | 65 | DS5
AAP05 | ANL Main body (AA01)Diluted detergent pump (P5) | 65 | DS3
AAP04 | ANL Main body (AA01)Detergent dilution water pump (P4) | 65 | DS2
AAP01 | ANL Main body (AA01)D.I water pump(P1) | 65 | DS1
AAFS02-4 | ANL Main body(AA01)Diluted detergent tank float SW (Empty) | 6E | DS8
AAFS02-3 | ANL Main body(AA01)Diluted detergent tank float SW (Lower) | 6E | DS7
AAFS02-2 | ANL Main body(AA01)Diluted detergent tank float SW (Upper) | 6E | DS6
AAFS02-1 | ANL Main body(AA01)Diluted detergent tank float SW (Overflow) | 6E | DS5
AAFS01-4 | ANL Main body(AA01)D.I. water tank float SW (Empty) | 6E | DS4
AAFS01-3 | ANL Main body(AA01)D.I. water tank float SW (Lower) | 6E | DS3
AAFS01-2 | ANL Main body(AA01)D.I. water tank float SW (Upper) | 6E | DS2
AAFS01-1 | ANL Main body(AA01)D.I. water tank float SW (Overflow) | 6E | DS1
AAFS06 | ANL Main body(AA01)Vacuum tank float SW | 6F | DS5
AAFS03-4 | ANL Main body(AA01)Detergent tank float SW (Empty) | 6F | DS4
AAFS03-3 | ANL Main body(AA01)Detergent tank float SW (Lower) | 6F | DS3
AAFS03-2 | ANL Main body(AA01)Detergent tank float SW (Upper) | 6F | DS2
AAFS03-1 | ANL Main body(AA01)Detergent tank float SW (Overflow) | 6F | DS1
AAFS04 | ANL Main body(AA01)Washing waste water tank float SW | 6F | DS2
AAFS05 | ANL Main body(AA01)Concentrated waste water tank float SW | 70 | DS1
SSR1 | ANL Main body(AA01)Status of D.I. water supply valve 1 (CAV01) | 74 | DS8
SSR2 | ANL Main body(AA01)Status of D.I water supply option valve(OPD) | 74 | DS7
ACD2 | ANL Main body(AA01) Status of D.I water supply valve 2 (V33) | 74 | DS1
AU5800 SMP
No. | Short name | ADD | Ds
CAS03 | Rack loading unit (CB01)/ Receiver movement Priority rack 1 rack detect sensor 2 | 05 | DS7
CAS02 | Rack loading unit (CB01)/ Receiver movement Priority rack 1 rack detect sensor 1 | 05 | DS6
CAS05 | Rack loading unit (CB01)/ Receiver movement CB1 Receiver stop position sensor | 05 | DS2
CAS04 | Rack loading unit (CB01)/ Receiver movement CB1 receiver origin sensor | 05 | DS1
CAS12 | Rack loading unit (CB02)/ Pusher movement CA unloading tray 2 detect sensor | 08 | DS4
CBS13 | Rack loading unit (CB02)/ Pusher movement CB2 pusher sensor | 08 | DS3
CBS11 | Rack loading unit (CB02)/ Pusher movement CB2 pusher origin sensor | 08 | DS1
CAS16 | Rack loading unit (CB02)/ Receiver movement Priority rack 1 rack detect sensor 2 | 0A | DS7
CAS15 | Rack loading unit (CB02)/ Receiver movement Priority rack 1 rack detect sensor 1 | 0A | DS6
CAS15 | Rack loading unit (CB02)/ Receiver movement CB2 Receiver stop position sensor | 0A | DS2
CAS14 | Rack loading unit (CB02)/ Receiver movement CB2 receiver origin sensor | 0A | DS1
CCS02 | ID read unit (CC01)/ Lever 1 movement CC lever 1 stop position sensor | 0C | DS8
CCS11 | ID read unit (CC01)/ Lever 1 movement Cup detect 5 | 0C | DS7
CCS10 | ID read unit (CC01)/ Lever 1 movement Cup detect 4 | 0C | DS6
CCS09 | ID read unit (CC01)/ Lever 1 movement Cup detect 3 | 0C | DS5
CCS08 | ID read unit (CC01)/ Lever 1 movement Cup detect 2 | 0C | DS4
CCS07 | ID read unit (CC01)/ Lever 1 movement Cup detect 1 | 0C | DS3
CAS17 | ID read unit (CC01)/ Lever 1 movement CA loading rack detect sensor | 0C | DS2
CCS01 | ID read unit (CC01)/ Lever 1 movement CC lever 1 origin sensor | 0C | DS1
CCS05 | ID read unit (CC01)/ Lever 2 movement Rack type | 10 | SD8-6
CCS06 | ID read unit (CC01)/ Lever 2 movement CC rack detect sensor | 10 | DS3
CCS04 | ID read unit (CC01)/ Lever 2 movement CC lever 2 stop position sensor | 10 | DS2
CCS03 | ID read unit (CC01)/ Lever 2 movement CC lever 2 origin sensor | 10 | DS1
CFS03 | Unloading rack feed unit(CF01)/ Lever movement CF lever stop position sensor 2 | 12 | DS3
CFS02 | Unloading rack feed unit(CF01)/ Lever movement CF lever stop position sensor 1 | 12 | DS2
CFS01 | Unloading rack feed unit(CF01)/ Lever movement CF lever origin sensor | 12 | DS1
CAS14 | Rack unloading unit (CG01)/ Lver movement CA unloading tray 2 detect sensor | 15 | DS7
CAS13 | Rack unloading unit (CG01)/ Lver movement CA unloading tray 1 detect sensor | 15 | DS6
CGS05 | Rack unloading unit (CG01)/ Lver movement CG rack detect sensor | 15 | DS5
CAS10 | Rack unloading unit (CG01)/ Lver movement CA unloading rack detect sensor 2 | 15 | DS4
CAS09 | Rack unloading unit (CG01)/ Lver movement CA unloading rack detect sensor 1 | 15 | DS3
CGS02 | Rack unloading unit (CG01)/ Lver movement CG lever stop position sensor | 15 | DS2
CGS01 | Rack unloading unit (CG01)/ Lver movement CG lever origin sensor | 15 | DS1
CAS06 | Passing lane feed unit(CE01)/ Passing lane lever movement CA passing lane rack detect sensor | 17 | DS3
CES02 | Passing lane feed unit(CE01)/ Passing lane lever movement CE passing lane lever stop position sensor | 17 | DS2
CES01 | Passing lane feed unit(CE01)/ Passing lane lever movement CE passing lane lever origin sensor | 17 | DS1
CAS07 | Passing lane feed unit(CE11)/ Passing lane lever movement CA dispensing lane rack detect sensor | 19 | DS3
CES12 | Passing lane feed unit(CE11)/ Passing lane lever movement CE dispensing lane lever stop position sensor | 19 | DS2
CES11 | Passing lane feed unit(CE11)/ Passing lane lever movement CE dispensing lane lever origin sensor | 19 | DS1
CAS08 | Recovery rack feed unit (CL01)/ Lever movement CA recovery lane rack detect sensor | 1B | DS3
CLS02 | Recovery rack feed unit (CL01)/ Lever movement CL lever stop position sensor | 1B | DS2
CLS01 | Recovery rack feed unit (CL01)/ Lever movement CL lever origin sensor | 1B | DS1
CMS10 | ISE rack transport unit(CM01)/ Dispensing lane belt movement Dispensing lane rack pass detect sensor | 1D | DS5
CMS09 | ISE rack transport unit(CM01)/ Dispensing lane belt movement Dispensing lane over step sensor | 1D | DS4
CMS08 | ISE rack transport unit(CM01)/ Dispensing lane belt movement Dispensing lane step sensor | 1D | DS3
CMS07 | ISE rack transport unit(CM01)/ Dispensing lane lever movement Dispensing lane lever stop position sensor | 10 | DS7
CMS06 | ISE rack transport unit(CM01)/ Dispensing lane lever movement Dispensing lane lever origin sensor | 10 | DS1
CMS07 | ISE rack transport unit(CM01)/ Dispensing lane lever movement Dispensing lane lever stop position sensor Latch | 21 | DS7
CMS11 | ISE rack recovery unit(CM11)/ Recovery lane belt movement Recovery lane rack present sensor | 24 | DS3
CMS05 | ISE rack transport unit(CM01)/ Passing lane belt movement Passing lane rack pass detect sensor | 27 | DS5
CMS04 | ISE rack transport unit(CM01)/ Passing lane belt movement Passing lane over step sensor | 27 | DS4
CMS03 | ISE rack transport unit(CM01)/ Passing lane belt movement Passing lane step sensor | 27 | DS3
CMS02 | ISE rack transport unit(CM01)/ Passing lane lever movement Passing lane lever stop position sensor | 2A | DS7
CMS01 | ISE rack transport unit(CM01)/ Passing lane lever movement Passing lane lever origin sensor | 2A | DS1
CMS02 | ISE rack transport unit(CM01)/ Passing lane lever movement Passing lane lever stop position sensor Latch | 2B | DS7
CHS05 | ANL1 rack transfer unit(CH01)/ Passing lane belt movement Passing lane rack pass detect sensor 1 | 2F | DS7
CHS02 | ANL1 rack transfer unit(CH01)/ Passing lane lever movement Passing lane lever stop position sensor | 32 | DS7
CHS04 | ANL1 rack transfer unit(CH01)/ Passing lane lever movement Passing lane over step sensor | 32 | DS4
CHS03 | ANL1 rack transfer unit(CH01)/ Passing lane lever movement Passing lane step sensor | 32 | DS3
CHS01 | ANL1 rack transfer unit(CH01)/ Passing lane lever movement Passing lane lever origin sensor | 32 | DS1
CHS13 | ANL1 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 rack present sensor | 39 | DS5
CHS12 | ANL1 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 error step sensor | 39 | DS4
CHS11 | ANL1 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 step sensor | 39 | DS3
CHS10 | ANL1 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 refuge stop position sensor | 39 | DS2
CHS09 | ANL1 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 refuge origin sensor | 39 | DS1
CHS13 | ANL1 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 rack present sensor Latch | 3A | DS5
CHS12 | ANL1 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 error step sensor Latch | 3A | DS4
CHS11 | ANL1 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 step sensor Latch | 3A | DS3
CHS08 | ANL1 rack transfer unit(CH01)/ Dispensing lane lever 1 movement Dispensing lane lever 1 stop position sensor | 3D | DS7
CHS07 | ANL1 rack transfer unit(CH01)/ Dispensing lane lever 1 movement Dispensing lane lever 1 origin sensor | 3D | DS1
CHS08 | ANL1 rack transfer unit(CH01)/ Dispensing lane lever 1 movement Dispensing lane lever 1 stop position sensor Latch | 3E | DS7
CHS22 | ANL1 rack transfer unit (CH01) / Dispensing lane lever 2 movement Dispensing lane post-dispensing stopper rack detect sensor | 42 | DS8
CHS15 | ANL1 rack transfer unit (CH01) / Dispensing lane lever 2 movement Dispensing lane lever 2 stop position sensor | 42 | DS7
CHS14 | ANL1 rack transfer unit (CH01) / Dispensing lane lever 2 movement Dispensing lane lever 2 origin sensor | 42 | DS1
CHS15 | ANL1 rack transfer unit (CH01)/ Dispensing lane lever 2 movement Dispensing lane lever 2 stop position sensor Latch | 43 | DS7
CHS19 | ANL1 rack transfer unit(CH01)/ Dispensing lane lever 2 refuge movement Dispensing lane lever 2 error step sensor | 46 | DS4
CHS18 | ANL1 rack transfer unit(CH01)/ Dispensing lane lever 2 refuge movement Dispensing lane lever 2 step sensor | 46 | DS3
CHS17 | ANL1 rack transfer unit(CH01)/ Dispensing lane lever 2 refuge movement Dispensing lane lever 2 refuge stop position sensor | 46 | DS2
CHS16 | ANL1 rack transfer unit(CH01)/ Dispensing lane lever 2 refuge movement Dispensing lane lever 2 refuge origin sensor | 46 | DS1
CHS21 | ANL1 rack transfer unit(CH01)/ Dispensing lane stopper movement Dispensing lane post-dispensing stopper stop position sensor | 48 | DS2
CHS20 | ANL1 rack transfer unit(CH01)/ Dispensing lane stopper movement Dispensing lane post-dispensing stopper origin sensor | 48 | DS1
CHS25 | ANL 1 rack recovery unit(CH11) /Recovery lane belt movement Recovery lane rack present sensor | 4A | DS8
CHS24 | ANL 1 rack recovery unit(CH11) /Recovery lane belt movement Recovery lane rack pass detect sensor | 4A | DS7
CHS06 | ANL1 rack transfer unit(CH01)/ Dispensing lane transfer belt movement Passing lane rack pass detect sensor 2 | 4C | DS7
CHS26 | ANL1 rack transfer unit(CH01)/ Dispensing lane transfer belt movement Dispensing lane rack pass detect sensor 2 | 4C | DS3
CKS04 | ANL 1 Lane change unit(CK01)/ Lane change movement Lane change guide sensor | 4E | DS4
CKS03 | ANL 1 Lane change unit(CK01)/ Lane change movement LLane change stop position sensor 2 (On recovery lane) | 4E | DS3
CKS02 | ANL 1 Lane change unit(CK01)/ Lane change movement Lane change stop position sensor 1 (On dispensing lane) | 4E | DS2
CKS01 | ANL 1 Lane change unit(CK01)/ Lane change movement Lane change origin sensor (On passing lane) | 4E | DS1
CKS06 | ANL 1 Lane change unit(CK01)/ Lane change stopper movement Lane change stopper stop position sensor | 50 | DS2
CKS05 | ANL 1 Lane change unit(CK01)/ Lane change stopper movement Lane change stopper origin sensor | 50 | DS1
CHS05 | ANL2 rack transfer unit(CH01)/ Passing lane belt movement Passing lane rack pass detect sensor 1 | 53 | DS7
CHS02 | ANL2 rack transfer unit(CH01)/ Passing lane lever movement Passing lane lever stop position sensor | 56 | DS7
CHS04 | ANL2 rack transfer unit(CH01)/ Passing lane lever movement Passing lane over step sensor | 56 | DS4
CHS03 | ANL2 rack transfer unit(CH01)/ Passing lane lever movement Passing lane step sensor | 56 | DS3
CHS01 | ANL2 rack transfer unit(CH01)/ Passing lane lever movement Passing lane lever origin sensor | 56 | DS1
CHS302 | ANL2 rack transfer unit(CH01)/ Passing lane lever movement Passing lane lever stop position sensor Latch | 57 | DS7
CHS23 | ANL2 rack transfer unit(CH01)/ Dispensing lane belt movement Dispensing lane rack pass detect sensor 1 | 5A | DS7
CHS13 | ANL2 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 rack present sensor | 5D | DS5
CHS12 | ANL2 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 error step sensor | 5D | DS4
CHS11 | ANL2 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 step sensor | 5D | DS3
CHS10 | ANL2 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 refuge stop position sensor | 5D | DS2
CHS09 | ANL2 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 refuge origin sensor | 5D | DS1
CHS313 | ANL2 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 rack present sensor Latch | 5E | DS5
CHS312 | ANL2 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 error step sensor Latch | 5E | DS4
CHS311 | ANL2 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 step sensor Latch | 5E | DS3
CHS08 | ANL2 rack transfer unit(CH01)/ Dispensing lane lever 1 movement Dispensing lane lever 1 stop position sensor | 61 | DS7
CHS07 | ANL2 rack transfer unit(CH01)/ Dispensing lane lever 1 movement Dispensing lane lever 1 origin sensor | 61 | DS1
CHS08 | ANL2 rack transfer unit(CH01)/ Dispensing lane lever 1 movement Dispensing lane lever 1 stop position sensor Latch | 62 | DS7
CHS22 | ANL2 rack transfer unit (CH01) / Dispensing lane lever 2 movement Dispensing lane post-dispensing stopper rack detect sensor | 66 | DS8
CHS15 | ANL2 rack transfer unit (CH01) / Dispensing lane lever 2 movement Dispensing lane lever 2 stop position sensor | 66 | DS7
CHS14 | ANL2 rack transfer unit (CH01) / Dispensing lane lever 2 movement Dispensing lane lever 2 origin sensor | 66 | DS1
CHS15 | ANL2 rack transfer unit (CH01)/ Dispensing lane lever 2 movement Dispensing lane lever 2 stop position sensor Latch | 67 | DS7
CHS19 | ANL2 rack transfer unit(CH01)/ Dispensing lane lever 2 refuge movement Dispensing lane lever 2 error step sensor | 6A | DS4
CHS18 | ANL2 rack transfer unit(CH01)/ Dispensing lane lever 2 refuge movement Dispensing lane lever 2 step sensor | 6A | DS3
CHS17 | ANL2 rack transfer unit(CH01)/ Dispensing lane lever 2 refuge movement Dispensing lane lever 2 refuge stop position sensor | 6A | DS2
CHS16 | ANL2 rack transfer unit(CH01)/ Dispensing lane lever 2 refuge movement Dispensing lane lever 2 refuge origin sensor | 6A | DS1
CHS21 | ANL2 rack transfer unit(CH01)/ Dispensing lane stopper movement Dispensing lane post-dispensing stopper stop position sensor | 6C | DS2
CHS20 | ANL2 rack transfer unit(CH01)/ Dispensing lane stopper movement Dispensing lane post-dispensing stopper origin sensor | 6C | DS1
CHS25 | ANL2 rack recovery unit(CH11) /Recovery lane belt movement Recovery lane rack present sensor | 6E | DS8
CHS24 | ANL2 rack recovery unit(CH11) /Recovery lane belt movement Recovery lane rack pass detect sensor | 6E | DS7
CHS06 | ANL2 rack transfer unit(CH01)/ Dispensing lane transfer belt movement Passing lane rack pass detect sensor 2 | 70 | DS7
CHS26 | ANL2 rack transfer unit(CH01)/ Dispensing lane transfer belt movement Dispensing lane rack pass detect sensor 2 | 70 | DS3
CKS04 | ANL 2 Lane change unit(CK01)/ Lane change movement Lane change guide sensor | 72 | DS4
CKS03 | ANL2Lane change unit(CK01)/ Lane change movement LLane change stop position sensor 2 (On recovery lane) | 72 | DS3
CKS02 | ANL2 Lane change unit(CK01)/ Lane change movement Lane change stop position sensor 1 (On dispensing lane) | 72 | DS2
CKS01 | ANL 2 Lane change unit(CK01)/ Lane change movement Lane change origin sensor (On passing lane) | 72 | DS1
CKS06 | ANL2 Lane change unit(CK01)/ Lane change stopper movement Lane change stopper stop position sensor | 74 | DS2
CKS05 | ANL 2 Lane change unit(CK01)/ Lane change stopper movement Lane change stopper origin sensor | 74 | DS1
CHS02 | ANL3 rack transfer unit(CH01)/ Passing lane lever movement Passing lane lever stop position sensor | 7A | DS7
CHS04 | ANL3 rack transfer unit(CH01)/ Passing lane lever movement Passing lane over step sensor | 7A | DS4
CHS03 | ANL3 rack transfer unit(CH01)/ Passing lane lever movement Passing lane step sensor | 7A | DS3
CHS01 | ANL3 rack transfer unit(CH01)/ Passing lane lever movement Passing lane lever origin sensor | 7A | DS1
CHS302 | ANL3 rack transfer unit(CH01)/ Passing lane lever movement Passing lane lever stop position sensor Latch | 7B | DS7
CHS23 | ANL3 rack transfer unit(CH01)/ Dispensing lane belt movement Dispensing lane rack pass detect sensor 1 | 7E | DS7
CHS13 | ANL3 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 rack present sensor | 81 | DS5
CHS12 | ANL3rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 error step sensor | 81 | DS4
CHS11 | ANL3 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 step sensor | 81 | DS3
CHS10 | ANL3 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 refuge stop position sensor | 81 | DS2
CHS09 | ANL3 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 refuge origin sensor | 81 | DS1
CHS313 | ANL3 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 rack present sensor Latch | 82 | DS5
CHS312 | ANL3 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 error step sensor Latch | 82 | DS4
CHS311 | ANL3 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 step sensor Latch | 82 | DS3
CHS08 | ANL3 rack transfer unit(CH01)/ Dispensing lane lever 1 movement Dispensing lane lever 1 stop position sensor | 85 | DS7
CHS07 | ANL3 rack transfer unit(CH01)/ Dispensing lane lever 1 movement Dispensing lane lever 1 origin sensor | 85 | DS1
CHS08 | ANL3 rack transfer unit(CH01)/ Dispensing lane lever 1 movement Dispensing lane lever 1 stop position sensor Latch | 86 | DS7
CHS22 | ANL3 rack transfer unit (CH01) / Dispensing lane lever 2 movement Dispensing lane post-dispensing stopper rack detect sensor | 8A | DS8
CHS15 | ANL3 rack transfer unit (CH01) / Dispensing lane lever 2 movement Dispensing lane lever 2 stop position sensor | 8A | DS7
CHS14 | ANL3 rack transfer unit (CH01) / Dispensing lane lever 2 movement Dispensing lane lever 2 origin sensor | 8A | DS1
CHS15 | ANL3 rack transfer unit (CH01)/ Dispensing lane lever 2 movement Dispensing lane lever 2 stop position sensor Latch | 8B | DS7
CHS19 | ANL3 rack transfer unit(CH01)/ Dispensing lane lever 2 refuge movement Dispensing lane lever 2 error step sensor | 8E | DS4
CHS18 | ANL3 rack transfer unit(CH01)/ Dispensing lane lever 2 refuge movement Dispensing lane lever 2 step sensor | 8E | DS3
CHS17 | ANL3 rack transfer unit(CH01)/ Dispensing lane lever 2 refuge movement Dispensing lane lever 2 refuge stop position sensor | 8E | DS2
CHS16 | ANL3 rack transfer unit(CH01)/ Dispensing lane lever 2 refuge movement Dispensing lane lever 2 refuge origin sensor | 8E | DS1
CHS21 | ANL3 rack transfer unit(CH01)/ Dispensing lane stopper movement Dispensing lane post-dispensing stopper stop position sensor | 90 | DS2
CHS20 | ANL3 rack transfer unit(CH01)/ Dispensing lane stopper movement Dispensing lane post-dispensing stopper origin sensor | 90 | DS1
CHS25 | ANL3 rack recovery unit(CH11) /Recovery lane belt movement Recovery lane rack present sensor | 92 | DS8
CHS24 | ANL3 rack recovery unit(CH11) /Recovery lane belt movement Recovery lane rack pass detect sensor | 92 | DS7
CHS06 | ANL3 rack transfer unit(CH01)/ Dispensing lane transfer belt movement Passing lane rack pass detect sensor 2 | 94 | DS7
CHS26 | ANL3 rack transfer unit(CH01)/ Dispensing lane transfer belt movement Dispensing lane rack pass detect sensor 2 | 94 | DS3
CKS04 | ANL3 Lane change unit(CK01)/ Lane change movement Lane change guide sensor | 96 | DS4
CKS03 | ANL3Lane change unit(CK01)/ Lane change movement LLane change stop position sensor 2 (On recovery lane) | 96 | DS3
CKS02 | ANL3 Lane change unit(CK01)/ Lane change movement Lane change stop position sensor 1 (On dispensing lane) | 96 | DS2
CKS01 | ANL 3 Lane change unit(CK01)/ Lane change movement Lane change origin sensor (On passing lane) | 96 | DS1
CKS06 | ANL3 Lane change unit(CK01)/ Lane change stopper movement Lane change stopper stop position sensor | 98 | DS2
CKS05 | ANL 3 Lane change unit(CK01)/ Lane change stopper movement Lane change stopper origin sensor | 98 | DS1
CHS02 | ANL4 rack transfer unit(CH01)/ Passing lane lever movement Passing lane lever stop position sensor | 9E | DS7
CHS04 | ANL4 rack transfer unit(CH01)/ Passing lane lever movement Passing lane over step sensor | 9E | DS4
CHS03 | ANL4 rack transfer unit(CH01)/ Passing lane lever movement Passing lane step sensor | 9E | DS3
CHS01 | ANL4 rack transfer unit(CH01)/ Passing lane lever movement Passing lane lever origin sensor | 9E | DS1
CHS302 | ANL4 rack transfer unit(CH01)/ Passing lane lever movement Passing lane lever stop position sensor Latch | 9F | DS7
CHS23 | ANL4 rack transfer unit(CH01)/ Dispensing lane belt movement Dispensing lane rack pass detect sensor 1 | A2 | DS7
CHS13 | ANL4 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 rack present sensor | A5 | DS5
CHS12 | ANL4rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 error step sensor | A5 | DS4
CHS11 | ANL4 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 step sensor | A5 | DS3
CHS10 | ANL4 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 refuge stop position sensor | A5 | DS2
CHS09 | ANL4 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 refuge origin sensor | A5 | DS1
CHS313 | ANL4 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 rack present sensor Latch | A6 | DS5
CHS312 | ANL4 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 error step sensor Latch | A6 | DS4
CHS311 | ANL4 rack transfer unit(CH01)/ Dispensing lane lever 1 refuge movement Dispensing lane lever 1 step sensor Latch | A6 | DS3
CHS08 | ANL4 rack transfer unit(CH01)/ Dispensing lane lever 1 movement Dispensing lane lever 1 stop position sensor | A9 | DS7
CHS07 | ANL4 rack transfer unit(CH01)/ Dispensing lane lever 1 movement Dispensing lane lever 1 origin sensor | A9 | DS1
CHS08 | ANL4 rack transfer unit(CH01)/ Dispensing lane lever 1 movement Dispensing lane lever 1 stop position sensor Latch | AA | DS7
CHS22 | ANL4 rack transfer unit (CH01) / Dispensing lane lever 2 movement Dispensing lane post-dispensing stopper rack detect sensor | E5 | DS8
CHS15 | ANL4 rack transfer unit (CH01) / Dispensing lane lever 2 movement Dispensing lane lever 2 stop position sensor | E5 | DS7
CHS14 | ANL4 rack transfer unit (CH01) / Dispensing lane lever 2 movement Dispensing lane lever 2 origin sensor | E5 | DS1
CHS15 | ANL4 rack transfer unit (CH01)/ Dispensing lane lever 2 movement Dispensing lane lever 2 stop position sensor Latch | E6 | DS7
CHS19 | ANL4rack transfer unit(CH01)/ Dispensing lane lever 2 refuge movement Dispensing lane lever 2 error step sensor | B2 | DS4
CHS18 | ANL4 rack transfer unit(CH01)/ Dispensing lane lever 2 refuge movement Dispensing lane lever 2 step sensor | B2 | DS3
CHS17 | ANL4 rack transfer unit(CH01)/ Dispensing lane lever 2 refuge movement Dispensing lane lever 2 refuge stop position sensor | B2 | DS2
CHS16 | ANL4rack transfer unit(CH01)/ Dispensing lane lever 2 refuge movement Dispensing lane lever 2 refuge origin sensor | B2 | DS1
CHS21 | ANL4 rack transfer unit(CH01)/ Dispensing lane stopper movement Dispensing lane post-dispensing stopper stop position sensor | B4 | DS2
CHS20 | ANL4 rack transfer unit(CH01)/ Dispensing lane stopper movement Dispensing lane post-dispensing stopper origin sensor | B4 | DS1
CHS25 | ANL4 rack recovery unit(CH11) /Recovery lane belt movement Recovery lane rack present sensor | B6 | DS8
CHS24 | ANL4 rack recovery unit(CH11) /Recovery lane belt movement Recovery lane rack pass detect sensor | B6 | DS7
CHS06 | ANL4 rack transfer unit(CH01)/ Dispensing lane transfer belt movement Passing lane rack pass detect sensor 2 | B8 | DS7
CHS26 | ANL4 rack transfer unit(CH01)/ Dispensing lane transfer belt movement Dispensing lane rack pass detect sensor 2 | B8 | DS3
CKS04 | ANL4 Lane change unit(CK01)/ Lane change movement Lane change guide sensor | BB | DS4
CKS03 | ANL4Lane change unit(CK01)/ Lane change movement LLane change stop position sensor 2 (On recovery lane) | BB | DS3
CKS02 | ANL4 Lane change unit(CK01)/ Lane change movement Lane change stop position sensor 1 (On dispensing lane) | BB | DS2
CKS01 | ANL4 Lane change unit(CK01)/ Lane change movement Lane change origin sensor (On passing lane) | BB | DS1
CKS06 | ANL4 Lane change unit(CK01)/ Lane change stopper movement Lane change stopper stop position sensor | BD | DS2
CKS05 | ANL4 Lane change unit(CK01)/ Lane change stopper movement Lane change stopper origin sensor | BD | DS1
CDS02 | Rack x axis transfer unit(CD01)/ X direction movement X-axis stop position sensor | DF | DS7
CDS07 | Rack x axis transfer unit(CD01)/ X direction movement Buffer side rack detect 2 sensor | DF | DS6
CDS06 | Rack x axis transfer unit(CD01)/ X direction movement Buffer side rack detect 1 sensor | DF | DS5
CDS08 | Rack x axis transfer unit(CD01)/ X direction movement Rack detect sensor | DF | DS4
CDS05 | Rack x axis transfer unit(CD01)/ X direction movement Set side rack detect sensor | DF | DS3
CDS01 | Rack x axis transfer unit(CD01)/ X direction movement X-axis origin sensor | DF | DS1
CDS02 | Rack x axis transfer unit(CD01)/ X direction movement X-axis stop position sensor Latch | C0 | DS7
CDS03 | Rack x axis transfer unit(CD01)/ Up down movement Up/down stop position sensor | C3 | DS2
CDS04 | Rack x axis transfer unit(CD01)/ Up down movement Up/down origin sensor | C3 | DS1
CDS12 | Rack Y axis transfer unit(CD11)/ Y direction movement Y-axis stop position sensor | C5 | DS7
CAS05 | Rack Y axis transfer unit(CD11)/ Y direction movement CA Buffer cover close sensor | C5 | DS4
CAS04 | Rack Y axis transfer unit(CD11)/ Y direction movement Priority rack cover close sensor | C5 | DS4
CDS11 | Rack Y axis transfer unit(CD11)/ Y direction movement Y-axis origin sensor | C5 | DS1
CDS12 | Rack Y axis transfer unit(CD11)/ Y direction movement Y-axis stop position sensor Latch | C6 | DS7
CALED01 | SMP Rack set DIAG LED Rack Set/DIAG SW & LED | C8 | DS3
CALED08 | SMP Rack set DIAG LED Rack Priority rack orange LED | C8 | DS1
AU5800 ISE
No. | Short name | ADD | Ds
IAS11 | S Dispenser unit (SA02)/ Up dowm movement Pressure sensor check | 03 | DS3
SAS22 | S Dispenser unit (SA02)/ Up dowm movement Lower point sensor | 03 | DS2
SAS21 | S Dispenser unit (SA02)/ Up dowm movement Upper point sensor | 03 | DS1
SAS32 | Sample inside wash dispenser (SA01) / Up down movement Lower point sensor | 05 | DS2
SAS31 | Sample inside wash dispenser (SA01) / Up down movement Upper point sensor | 05 | DS1
FAS30 | ISE Sample Transfer unit (FA21)/ Up down movement Sample level detection sensor | 08 | DS4
FAS29 | ISE Sample Transfer unit (FA21)/ Up down movement Sample probe crash detection sensor | 08 | DS3
FAS21 | ISE Sample Transfer unit (FA21)/ Up down movement Upper position sensor | 08 | DS2
FAS30 | ISE Sample Transfer unit(FA21)/ Up down movement Sample level detection sensor Latch | 09 | DS4
FAS29 | ISE Sample Transfer unit(FA21)/ Up down movement Sample probe crash detection sensor Latch | 09 | DS3
FAS27 | ISE Sample Transfer unit(FA21)/ Rotation movement Rotation stop position sensor | 0B | DS3
FAS25 | ISE Sample Transfer unit(FA21)/ Rotation movement Rotation stop position sensor | 0B | DS2
FAS24 | ISE Sample Transfer unit(FA21)/ Rotation movement Origin sensor | 0B | DS1
IAS02 | ISE 1 Bottle / Cover Status MID bottle detection sensor | 0E | DS3
IAS03 | ISE 1 Bottle / Cover Status Buffer bottle detection sensor | 0E | DS2
IAS09 | ISE 1 Bottle / Cover Status Cover open reed SW | 0E | DS1
IAS01 | ISE 1 Buffer Dispenser 1 (SA01)/ Up down movement REF b bottle detection sensor 1 | 11 | DS3
SAS02 | ISE 1 Buffer Dispenser 1 (SA01)/ Up down movement Lower point sensor | 11 | DS2
SAS01 | ISE 1 Buffer Dispenser 1 (SA01)/ Up down movement Upper point sensor | 11 | DS1
IAS06 | ISE 2 Bottle Status MID bottle detection sensor 2 | 14 | DS3
IAS07 | ISE 2 Bottle Status Buffer bottle detection sensor 2 | 14 | DS22
IAS05 | Buffer Dispenser 2(SA01)/ Up down movement REF b bottle detection sensor 2 | 17 | DS3
SAS12 | Buffer Dispenser 2(SA01)/ Up down movement Lower point sensor | 17 | DS2
SAS11 | Buffer Dispenser 2(SA01)/ Up down movement Upper point sensor | 17 | DS1
IAV03 | ISE Main body （CPU） S probe wash water refresh valve | 19 | DS3
IAV02 | ISE Main body （CPU） S probe outer wash valve | 19 | DS2
IAV01 | ISE Main body （CPU） S probe wash inner valve | 19 | DS1
IASW02 | ISE Main body （CPU） DIAG_SW | 1A | DS3
IASW04 | ISE Main body （CPU） STOP_SW | 1A | DS2
IASW03 | ISE Main body （CPU）StandBy_SW | 1A | DS1
IAS04 | ISE 1 Main Body (IA11) Liquid level detector 1 | 20 | DS8
IAS12 | ISE 1 Main Body (IA11) Cable disconnect 1 | 20 | DS5
IAS08 | ISE 2 Main Body (IA11) Liquid level detector 2 | 21 | DS8
IAS13 | ISE 2 Main Body (IA11) Cable disconnect 2 | 21 | DS5